## Version
Current version and date:
Version 1.0.4 27/05/2011
Fixed problem with relative CTDI for Siemens Definition AS
Added narrow collimation data for Philips Brilliance 40 and 64
Version history:
Version 1.0.3 24/8/2010
Added Philips Brilliance 40 and 64 scanners
Corrected errors in Emotion 6 head CTDI centre values at 80 and 110 kV
Version 1.0.2 12/11/2009
Changed VBA code to avoid Excel 2007 crash: Phantom worksheet unprotected before changing the size/position of the rectangle indicating the scan range
Version 1.0.1 04/11/2009
Modified VBA code for phantom range calculation to avoid crash in Excel 2007 on some machines
Version 1.0 28/08/2009
Added ability to calculate effective dose according to ICRP 60 or ICRP 103, incorporating new weighted and remainder organs and weighting factors
Removed patient sex selection box. All doses now listed for hermaphrodite phantom, with sex dependent organs listed separately
Added eye lenses to phantom diagram, and lens dose to ScanCalculation worksheet
Changed calculation of colon dose from LLI to mass weighted mean of ULI and LLI, and removed ULI from ICRP60 remainder organs, as recommended by ICRP60
Changed dose to prostate from LLI to bladder (Prostate not in SR250, bladder is closer match)
Added Philips Big Bore and Siemens Definition AS
Corrected CTDIair values for Toshiba Aquilion (were quoted for 16, not 8 mm collimation)
Finally bumped version number to 1.0!
Version 0.99x 20/01/2006
Added LightSpeed RT (meant for 0.99w, but data was under review and removed before publication)
Added LightSpeed VCT
Version 0.99w 16/05/2005
Fixed problem when small ranges are used (< 5 mm), which start and end at non integer slab positions within a single slab
Added GE LightSpeed Pro 16, LightSpeed RT, Siemens Sensation Open and Sensation 64
Added Philips Brilliance 16 and Brilliance 16 Power to description of Mx8000 IDT data
Updated links to NRPB SR250, which is now part of the HPA (www.hpa.org.uk)
Updated some of the text in the Introduction to reflect changes over recent versions.
Version 0.99v 17/06/2004
Changed nCTDIw cell to be editable, and added a 'look up' button
Added Siemens Emotion 6 and Sensation 10
Added CTDI information for Siemens Emotion Duo at 80 and 110kV
Improved MCSET load routine to avoid odd errors. Thanks to Ed McDonagh, Royal Marsden
Version 0.99u 12/12/2003
Fixed error for scanners with more than 10 collimation options, where incorrect relative CTDI was chosen.
Fixed error with international locations that use '.' as a thousands separator
Added Philips Aura CTDI data and 135kV data for GE Sytec
Version 0.99t 24/7/2003
Relative CTDI at 1 x 1.25mm corrected for the LightSpeed Plus
Added CTDI data for Toshiba Aquilion 16 and Philips Mx8000 IDT
Modified collimation data for GE NX/i / ZX/i and Philips AV to allow correct selection
Version 0.99s 07/05/2003
Corrected errors in Mx8000 data. Head phantom data and body 140kV data updated.
Version 0.99r 02/04/2003
Added 'Paediatric' worksheet, giving information on doses to children relative to adults
Separated organ and effective dose calculations to the new 'DoseCalculations' worksheet
Improved rounding of results so that only two significant figures are displayed for calculated organ and effective doses.
Version 0.99q 20/02/2003
Added CTDI data for Toshiba Asteion Multi, pre and post 2002, and Toshiba Asteion Dual.
Duplicated Marconi Mx8000 and AcQSim data so they are also listed as Philips scanners.
Version 0.99p 17/02/2003
Removed typographical error on 'Collimation' worksheet that stops correct selection of CTDI at 'different collimations for the LightSpeed Ultra and LightSpeed 16.
Version 0.99o 28/01/2003
Fixed error on Excel 97, which crashed with decimal separator code introduced in 0.99m
Version 0.99n 17/01/2003
Included CTDI data for Siemens Sensation 4, Sensation 16, GE LightSpeed Ultra and LightSpeed 16.
Changed Visual Basic code to avoid use of absolute references to cells on selections worksheet
Version 0.99m 1/07/2002
Fixed problem with computers with comma as decimal separator having problems importing MC data
Moved organ dose calculation data from ScanCalculation to Selections
Included a cell showing CTDIw in ScanCalculation
Updated nCTDIw calculation to include relative CTDI at different collimations
Version 0.99l 17/06/2002
Added data and selector for relative CTDIs at different collimations
Updated GE LightSpeed match to include LightSpeed Plus
Added CTDI data for Siemens Emotion Duo
Version 0.99k 24/01/2002
Incorporated CTDI values for Philips AcQsim scanner.
Updated CTDI values for GE LightSpeed, and Toshiba Aquilion Multi
Version 0.99j 23/10/2001
Minor fix to recognise data sets 301-323 as being loaded (not relevant for most users)
Version 0.99i 19/10/2001
Added calculation of CTDIw and DLP values
Included scanner matching data for GE ZX/i and NX/i
Modified ScanCalculation!C11 to account for early versions of SR250 that number the data sets 301-323 rather than 1-23.
Version 0.99h 30/4/2001
Changed mAs/slice to mAs/rotation in ScanCalculation
Updated LightSpeed and Aquilion results to use correct CTDIair for Body filter
Version 0.99g 20/2/2001
Corrected definition of pitch in Introduction.
Added 'Scan Description / Comments' box to ScanCalculation
Version 0.99f 15/2/2001
Removed 'actual' kV setting from ScanCalculation, as match and actual kV should be identical
Changed remainder organ calculation so that organ is included if the dose is greater than or equal to the maximum dose to listed radiosensitive organs
Pointed out the potential problem with decimal points in 'Introduction'
Version 0.99e 10/10/2000
Introduced version worksheet to detail changes in different versions of CTDosimetry.xls
Added information for Siemens Volume Access and Volume Zoom
Fixed CTDI lookup formula to correctly look up Head CTDI in air
Updated web links to new ImPACT web site at www.impactscan.org

## Introduction
ImPACT CT Patient Dosimetry Calculator
Version 1.0.4 27/05/2011
Introduction
This spreadsheet is a tool for calculating patient organ and effective doses from CT scanner examinations. It makes use of the NRPB Monte Carlo dose data sets produced in report SR250 (link at bottom of page). SR250 provides normalised organ dose data for irradiation of a mathematical phantom by a range of CT scanners.
As SR250 was produced in 1993, it does not include data for more modern scanners. To overcome this problem, the ImPACT CT scanner dose survey was carried out by physicists in the UK and Europe.
This work, provides a method for 'matching' the dose distribution of newer scanners to scanners included in SR250. The matching results are included in this spreadsheet. As new scanners are introduced, their matches will be included in updates to this spreadsheet. More details can be found on the dose survey page on the ImPACT website (link below)
The results produced by the CTDosimetry spreadsheets have been checked against those produced by CTDOSE, produced by John Le Heron, the standard software used to calculate doses from the NRPB SR250 datasets. The two methods produce identical results for a range of scans and scanners, with the exception of small differences between the doses calculated for muscle and for the 'remainder' organs. These differences are present despite an apparent similarity in calculation method, and are typically 1-2%.
Installation
The system should work on any PC with Microsoft Excel 2000 or above. It has not yet been tested on a Apple computer, but it is anticipated that it should work on a Mac.
Installation is fairly simple, and only requires the SR250 data sets (MCSET01.DAT to MCSET23.DAT) to be present in the same directory as this spreadsheet. (SR250 is sold by the NRPB - see link below)
Macros are used on this spreadsheet for a variety of purposes. Depending on your version of Excel , and macro options, the security level may have to be switched to 'medium' (select 'Tools' -> 'Options' -> 'Security' -> 'Macro Security'), and/or ImPACT added to your trusted Macro sources.
This spreadsheet has been checked for macro viruses, and the logic and calculations been tested extensively, however ImPACT accept no responsibility for loss or damage incurred as a result of its use.
Worksheets
CTDosimetry.xls consists of 12 worksheets
Introduction | Provides an introduction and instructions for use
Scan Calculation | The data entry and results sheet
Paediatric | Information on relative doses to adult and paediatric patients
Phantom | Allows interactive selection of the scan range used for dose calculation using a diagram of the phantom used to generate SR250
Scanners | Provides data on CT scanner models, including CTDI in air and phantom, as well as the scanner matching data
MatchData | Gives data required to perform the scanner matchings in the Scanners worksheet
Collimation | Lists relative CTDI values at different collimations for a range of CT scanners. These values are more useful for multi-slice scanners, as the CTDI can vary considerably over the range of available collimations
MonteCarloData | Contains the unformatted SR250 data set.
Doses | Contains the formatted dose data from the SR250 data set that is currently loaded.
DoseCalculations | Performs the organ dose calculations, and calculation of remainder organ doses etc.
Selections | Provides data for the drop down selection boxes in the ScanCalculation worksheet, and performs calculations for 'remainder' organ doses
Version | Details changes made in each version, from version 0.99e onwards
There are also a number of Visual Basic macros used by CTDosimetry.xls, held in the modules PhantomDiagram, ScannerSelection and UpdateDataSet
Using CTDosimetry.xls
To calculate doses using CTDosimetry.xls, the user must enter a number of parameters relating to the scanner and the scan series
The following four selections, made in the top left box on the ScanCalculations worksheet define the Monte Carlo data set that is used:
Manufacturer | Select the scanner manufacturer from the drop down list
Scanner | Select the scanner model or scanner model group for the drop down list.
kV | Choose the appropriate scan kV.
Scan Region | Choose head or body.
The Monte Carlo data set that is used for this combination of scanner, kV and body part is displayed in the cell marked 'Data Set'. The data set that is currently loaded is displayed below. If these do not match, no dose is calculated. To load the appropriate data set, and enable dose calculation, press the 'Update Data Set' button.
Scan and patient data is entered in the box on the top right of the ScanCalculations worksheet.
Tube current | The x-ray tube current. Note that this should be the actual scanner mA, and not the 'effective mAs' displayed on some multi-slice scanners
Rotation time | The scanner tube rotation time
Spiral pitch | The scanning pitch (table travel per rotation/total collimated slice width). For axial scanning, (couch increment)/(collimated slice width) should be used
mAs/rotation | The total mAs per gantry rotation. Do not enter data in this box - it is calculated automatically.
Effective mAs | The mAs/per rotation divided by the spiral pitch. This is a calculated value that provides a basis for comparison of spiral protocols with different pitches
Collimation | The total nominal x-ray beam width along the z-axis, selected from a range of possible values in the drop down box. This determines the relative CTDI compared to the reference (usually 10 mm) collimaiton.
Rel. CTDI | The CTDI at the selected collimated x-ray beam thickness, relative to the CTDI at the reference collimation (usually 10 mm)
CTDI (air) | The free in air CTDI100 value (in mGy/100mAs), as defined in EUR 16262: European Guidelines on Quality Criteria for Computed Tomography, pub. European Comission (link to this document at bottom of page). CTDI values for most of the scanners are listed on the Scanner Worksheet. Pressing the 'Look up' button will enter the value in this cell. The value in this cell is corrected for the relative CTDI value in the cell above.
CTDI (soft tissue) | The CTDI to ICRU muscle, used as an approximation to the dose to soft tissue within the body. This is the CTDI(air) x 1.07 for CT scanner energies
nCTDIw | Weighted CTDI measured in a standard CTDI phantom (normalised for 100 mAs). CTDIw = (CTDIcentre + 2 * CTDIperiphery)/3. See EUR16262 for more details (link below)
CTDIw | Weighted CTDI measured in a standard CTDI phantom. CTDIw = (CTDIcentre + 2 * CTDIperiphery)/3. See EUR16262 for more details (link below)
CTDIvol | Volume weighted CTDI, given by CTDIvol = CTDIw / Spiral Pitch
DLP | Dose Length Product, given by CTDIvol x Scan length
Start Position | The start position of the scan series. The diagram on the Phantom worksheet shows the position of the phantom's organs relative to the number scale, which is 0 at the base of the trunk. This value can be entered manually in the worksheet, or can be taken from the shaded area on the Phantom worksheet diagram. This can be adjusted using the up and down arrows. Pressing the 'Get From Phantom Diagram' button enters these values into the start and end position boxes in Scan Calculation.
End Position | The end position of the scan series - Note that this should include the slice thickness, so, for example, a single 5mm slice 20cm from the base of the trunk would have a start position of 20, and an end position of 20.5cm. Start and End position values are interchangeable
Organ weighting scheme | The calculation of effective dose is governed by the weighting of doses to individual organs according to tissue weighting factors given in ICRP publications 60 and 103. Changing thisvalue alters the organ weighting factors and organs used in effective dose calculation.
When the above values are entered, the doses to each of the individual organs, as defined by the SR250 data set appear in the cells below the scan parameters. These are combined according to the tissue weighting factors given in ICRP publications 60 or 103, to calculate an effective dose
In addition, the weighted CTDI (CTIDw), volume CTDIw (CTDIvol) and dose length product (DLP) are also displayed.
Note that not all of the organs listed in ICRP publication 60 and 103 are included in NRPB SR250. In order to approximate dose to the oesophagus, salivary gland and prostate, the thymus, brain and blsdder doses are used. The dose for muscle is approximated from the total body dose - dose to all other organs and contents. Three organ doses included in the ICRP 103 definition of remainder organ; extrathoracic region, lymphatic nodes and oral mucosa, are approximated by thyroid, muscle and brain.
Any questions?
This spreadsheet is updated with data on recent scanners as they are tested by ImPACT, so check back on the CTDosimetry page for updates. Also, if you have any problems with, or comments about this spreadsheet, please let us know.
Copyright
This spreadsheet may be distributed freely but remains copyright ImPACT. If you had this spreadsheet passed on to you by another user, please fill out the form on the CTDosimetry page of the ImPACT web site (link below). We would like to be able to keep track of how many people are using it.
Links
NRPB SR250
ImPACT Home Page
ImPACT CTDosimetry page
ImPACT Dose Survey and Scanner Matching Page
ImPACT Dose Survey
EUR 16262: European Guidelines on Quality Criteria for Computed Tomography
Contacting ImPACT:
Phone | +44 (0) 20 8725 3366
Fax | +44 (0) 20 8725 3969
email | impact@impactscan.org
ImPACT
Medical Physics Department
Knightsbridge Wing
St George's Hospital
Tooting
London SW17 0QT

## ScanCalculation
ImPACT CT Patient Dosimetry Calculator
Version 1.0.4 27/05/2011
Scanner Model: |  |  |  |  | Acquisition Parameters:
Manufacturer: |  |  |  |  | Tube current |  | 30 | mA
Scanner: |  |  |  |  | Rotation time |  | 1 | s
kV: |  |  |  |  | Spiral pitch |  | 1
Scan Region: |  |  |  |  | mAs / Rotation |  | 30 | mAs
Data Set | #N/A |  |  |  | Effective mAs |  | 30 | mAs
Current Data | MCSET00 |  |  |  | Collimation |  |  | mm
Scan range |  |  |  |  | Rel. CTDI | 1.4476 | 1.4476 | at selected collimation
Start Position | 42.5 | cm |  |  | CTDI (air) | #N/A | #N/A | mGy/100mAs
End Position | 64 | cm |  |  | CTDI (soft tissue) |  | #N/A | mGy/100mAs
[5 empty] | nCTDIw | #N/A | #N/A | mGy/100mAs
Organ weighting scheme
[5 empty] | CTDIw |  | #N/A | mGy
[5 empty] | CTDIvol |  | #N/A | mGy
[5 empty] | DLP |  | #N/A | mGy.cm
Organ |  | wT | HT (mGy) | wT.HT |  | Remainder Organs |  |  | HT (mGy)
Gonads |  | 0.2 | #N/A | #N/A |  | Adrenals |  |  | #N/A
Bone Marrow |  | 0.12 | #N/A | #N/A |  | Small Intestine |  |  | #N/A
Colon |  | 0.12 | #N/A | #N/A |  | Kidney |  |  | #N/A
Lung |  | 0.12 | #N/A | #N/A |  | Pancreas |  |  | #N/A
Stomach |  | 0.12 | #N/A | #N/A |  | Spleen |  |  | #N/A
Bladder |  | 0.05 | #N/A | #N/A |  | Thymus |  |  | #N/A
Breast |  | 0.05 | #N/A | #N/A |  | Uterus |  |  | #N/A
Liver |  | 0.05 | #N/A | #N/A |  | Muscle |  |  | #N/A
Oesophagus (Thymus) |  | 0.05 | #N/A | #N/A |  | Brain |  |  | #N/A
Thyroid |  | 0.05 | #N/A | #N/A |  | Not Applicable |  |  | N/A
Skin |  | 0.01 | #N/A | #N/A |  | Not Applicable |  |  | N/A
Bone Surface |  | 0.01 | #N/A | #N/A |  | Not Applicable |  |  | N/A
Not Applicable |  | 0 | 0 | #N/A |  | Not Applicable |  |  | N/A
Not Applicable |  | 0 | 0 | #N/A |  | Other organs of interest |  |  | HT (mGy)
Remainder |  | #N/A | #N/A | #N/A |  | Eye lenses |  |  | #N/A
#N/A |  | #N/A | #N/A | #N/A |  | Testes |  |  | #N/A
|  |  | Total Effective Dose (mSv) | #N/A |  | Ovaries |  |  | #N/A
[6 empty] | Uterus |  |  | #N/A
[6 empty] | Prostate |  |  | #N/A
Scan Description / Comments
© Nicholas Keat for ImPACT, 2000-2011
Imaging Performance Assessment of CT Scanners, an MHRA Evaluation centre
http://www.impactscan.org
Note on B8: Choose the closest kV to that being used in the scan
Note on B10: NRPB SR250 Data Set Number for this combination of Manufacturer, Scanner, kV and Scan Region. See http://tinyurl.com/SR250 to obtain this data
Note on B11: Data set currently loaded. If this is not the same as the 'data set', above, no dose will be given. Press update button to load new data
Note on C13: Scan start position, as defined by Phantom worksheet. (0cm is level with base of trunk, scans go caudo-cranially)
Note on C14: Scan End Position as defined in Phantom worksheet
Note on C34: The organs in cells C34 and C35 (Brain and Salivary Glands) are only included in ICRP 103 calculations
Note on C37: This organ is only present if using ICRP60 organ weighting scheme and the maximum remainder organ dose is greater than the maximum dose to an organ with a specified weight.
Note on D21: Organ Tissue weighting factor according to ICRP organ weighting scheme specified above
Note on E21: Organ Dose (mGy) quoted to two significant figures
Note on F21: Contribution to effective dose from each organ (Organ dose x Tissue weighting factor) quoted to two significant digits
Note on G11: X-ray beam width, used to look up relative CTDI. For multi-slice scanners, this is the total beam width. When #N/A is displayed, no collimation data is available for that scanner
Note on H21: Organs not given a tissue weighting factor, but listed in ICRP60 or 103 as potentially sensitive organs
Note on I6: Tube current setting
Note on I7: Tube rotation time
Note on I8: Pitch =Table travel/x-ray beam width. For multi-slice, use total beam width, not data acquisition or image width. Pitches are generally from 0.25 -> 2. If you are using pitches of 3 -> 8, the calculation will be incorrect! Non contiguous axial scanning can be accounted for here by using the ratio of irradiated length to total scan length (=1/packing factor)
Note on I9: Calculated tube mAs per rotation
Note on I10: Effective mAs, which is mAs per rotation divided by spiral pitch
Note on I12: CTDI at the above collimation, relative to the CTDI at 10mm collimation. Enter manually, or press the 'Look up' button to enable automatic look up. Manually entered values that are not equal to the looked up value are highlighted with red text
Note on I13: MEASURED CTDI in air at isocentre, per 100mAs, at the above collimation. Enter manually, or press the 'Look up' button to enable automatic look up. Manually entered values that are not equal to the looked up value are highlighted with red text
Note on I14: CTDI to ICRU muscle (CTDI to air x 1.07)
Note on I15: Weighted CTDI in Perspex CTDI phantom, normalised per 100mAs. Enter manually, or press the 'Look up' button to enable automatic look up. Manually entered values that are not equal to the looked up value are highlighted with red text.
Note on I17: Weighted CTDI in Perspex phantom
Note on I18: Weighted CTDI in Perspex, adjusted for mAs and Pitch
Note on I19: Dose Length Product (=CTDI x Scan Length)
Note on J30: This remainder organ depends on the ICRP weighting scheme used. Under ICRP60 it is Brain, under ICRP103, it is Gall Bladder
Note on J31: The heart is shown here for ICRP103 calculations only
Note on J32: The following three cells are used in ICRP103 calculations only, and include the extrathoracic region, lymph nodes and oral mucosa. As none of these organs are present in SR250, the most suitable organs have been used to provide an approximation to the relevent organ dose.
Note on J36: The lenses are of interest for deteministic reasons (cataract formation), rather than for effective dose calculations
Note on J37: The presence of these organs is dependent on the sex of the patient, and they are generally averaged for population based Effective Dose calculations. They are listed here for protocol optimisation purposes
Update Data Set
Look up
Get From Phantom Diagram
Look up
Look up

## Phantom
Start: | +1 | +10 | End: | +1 | +10 | 106 | 200.5
42.5 | -1 | -10 | 64 | -1 | -10 | 43 | 64.5
90
-10
0
10
80
70
60
50
40
30
20
Zoom In
Zoom Out

## Paediatric
Effective doses to paediatric patients
The phantom used to produce the Monte Carlo data sets in NRPB report SR250 is based upon a mathematical representation of an 'average' adult, but does not address the issue of dose to paediatric patients.
In the paper 'Influence of patient age on normalized effective doses calculated for CT examinations' (full reference below), the NRPB made Monte Carlo calculations on a range of phantoms corresponding to Newborn, 1 year old, 5 year old, 10 year old, 15 year old and adult patients.
This enabled the calculation of relative effective doses to patients of different ages from CT examinations for different body parts. Note that when doses are compared, the scan range is defined by the extent of a relevant organ, rather than by a set number of slices.
The calculations were carried out for three scanners, Siemens DRH, GE 9800 and Philips LX, all of which were modelled in the original SR250 report. These three scanners were chosen as they represent a wide range of scanner designs, in terms of geometry and beam filtration.
A summary of the results is presented below, showing the range of normalised effective doses at each age, relative to the adult dose. The range represents the minimum and maximum relative dose from the three scanners simulated.
|  | Head and Neck | Chest | Abdo & Pelvis
| Adult | 1 | 1 | 1
| 15 y | 1.1 | 1.0 - 1.1 | 1.0 - 1.1
| 10 y | 1.2 - 1.3 | 1.1 - 1.4 | 1.2 - 1.5
| 5 y | 1.6 - 1.7 | 1.2 - 1.6 | 1.2 - 1.6
| 1 y | 2.2 | 1.3 - 1.9 | 1.3 - 2.0
| Newborn (0 y) | 2.3 - 2.6 | 1.4 - 2.2 | 1.4 - 2.4
Typical normalised effective doses to paediatric patients relative to adults
Note that these figures are supplied to give an indication of relative adult and paediatric doses. They can be used to multiply the adult effective doses calculated by this spreadsheet to provide a broad estimate of the corresponding doses to children under similar conditions of exposure.
Reference:
Khursheed A, Hillier MC, Shrimpton PC and Wall BF. Influence of patient age on normalized effective doses calculated for CT examinations. Br J Radiol 2002; 75:819-830

## Scanners
Scanner | kVp | Sub-group | Scanner | CTDI (Head, mGy/100mAs) |  |  | CTDI (Body, mGy/100mAs) |  |  | ImPACT Factor |  |  |  | Scanner Match
Group |  |  |  | Air | Centre | Perip | Air | Centre | Perip | Head | Body |  |  | Head | Body
CG.a | 130 | CG.a.130 | CGR CE 10000,12000 | [6 empty] | 1.0827 | 1.0559 | 19 | 19 | 12 | 12
EL.a | 120 | EL.a.120 | Elscint Exel 2400 Elect | 18.776 | 13.1795 | 14.5466 | 18.5037 | 3.4805 | 6.232 | 1.0311 | 0.7971 | 16 | 11 | 19 | 20
EL.a | 140 | EL.a.140 | Elscint Exel 2400 Elect |  |  |  | 25.793 | 5.2187 | 8.587 |  | 0.84566 |  | 11 |  | 20
EL.b | 120 | EL.b.120 | Elscint CT Twin, Helicat | 18.56 | 12.9468 | 13.8592 | 19 | 3.82 | 6.49 | 1.0064 | 0.84656 | 14 | 11 | 18 | 20
GE.a | 120 | GE.a.120 | GE 8800 / 9000 Series | 14.0717 | 6.4557 | 6.0884 | 14.5664 | 2.0859 | 4.2149 | 0.64065 | 0.60617 | 5 | 7 | 6 | 7
GE.b | 120 | GE.b.120 | GE 9800 Series | 25.96 | 14.1 | 14.9 | 25.96 | 3.94 | 7.4 | 0.79429 | 0.63416 | 8 | 8 | 8 | 8
GE.b | 140 | GE.b.140 | GE 9800 Series | 34.11 | 19.39 | 19.97 | 34.11 | 5.67 | 9.95 | 0.81554 | 0.69018 | 9 | 10 | 9 | 9
GE.c | 80 | GE.c.080 | GE HiLight, HiSpeed, CT/i (No SmB) | 8.5438 | 4.2035 | 4.5009 | 8.52 | 1 | 1.9 | 0.73212 | 0.4716 | 6 | 5 | 10 | 10
GE.c | 100 | GE.c.100 | GE HiLight, HiSpeed, CT/i (No SmB) | 14 | 8.2 | 8.3 | 14 | 2.2 | 4.3 | 0.82966 | 0.66739 | 9 | 9 | 9 | 11
GE.c | 120 | GE.c.120 | GE HiLight, HiSpeed, CT/i (No SmB) | 19.3213 | 11.4281 | 11.9396 | 18.7652 | 3.2147 | 6.1309 | 0.85258 | 0.73056 | 9 | 11 | 13 | 5
GE.c | 140 | GE.c.140 | GE HiLight, HiSpeed, CT/i (No SmB) | 27.0106 | 16.8387 | 17.2181 | 25.7588 | 4.8266 | 8.9605 | 0.88341 | 0.80153 | 10 | 11 | 13 | 20
GE.d | 80 | GE.d.080 | GE HiSpeed CT/i with SmartBeam | 8.5438 | 4.2035 | 4.5009 | 8.4679 | 1.0366 | 2.7356 | 0.73212 | 0.553 | 6 | 6 | 10 | 10
GE.d | 100 | GE.d.100 | GE HiSpeed CT/i with SmartBeam | 14 | 8.2 | 8.3 | 13.86 | 2.5 | 5.6 | 0.82966 | 0.81198 | 9 | 11 | 9 | 20
GE.d | 120 | GE.d.120 | GE HiSpeed CT/i with SmartBeam | 19.3213 | 11.4281 | 11.9396 | 20.4347 | 3.8433 | 7.3397 | 0.85258 | 0.81118 | 9 | 11 | 13 | 20
GE.d | 140 | GE.d.140 | GE HiSpeed CT/i with SmartBeam | 27.0106 | 16.8387 | 17.2181 | 27.6945 | 5.7525 | 10.6551 | 0.88341 | 0.89774 | 10 | 13 | 13 | 18
GE.e | 120 | GE.e.120 | GE Max | 38.3695 | 18.7625 | 17.6551 | 38.3695 | 4.9828 | 8.7554 | 0.67706 | 0.51965 | 6 | 6 | 10 | 10
GE.f | 80 | GE.f.080 | GE Pace, Sytec | 20.05 | 8.0167 | 9.96 | 19.22 | 1.93 | 5.3 | 0.66428 | 0.44421 | 5 | 5 | 10 | 10
GE.f | 120 | GE.f.120 | GE Pace, Sytec | 41.5582 | 22.0763 | 23.8175 | 41.0371 | 6.2962 | 13.2268 | 0.78796 | 0.66367 | 7 | 9 | 11 | 11
GE.f | 135 | GE.f.135 | GE Pace, Sytec | 50.701 | 28.2577 | 30.0101 | 50.0652 | 8.374 | 16.2442 | 0.81545 | 0.71462 | 8 | 11 | 9 | 5
GE.f | 140 | GE.f.140 | GE Pace, Sytec | 55.5132 | 30.6764 | 33.4709 | 54.0598 | 9.4915 | 18.5396 | 0.82208 | 0.75611 | 9 | 11 | 9 | 5
GE.g | 120 | GE.g.120 | GE Prospeed | 36.6 | 20.7 | 22.3 | 36.6 | 5.7 | 11.8 | 0.83334 | 0.67201 | 9 | 9 | 9 | 11
GE.g | 140 | GE.g.140 | GE Prospeed | 47.2 | 28.4 | 30.1 | 47.2 | 8.3 | 15.9 | 0.87332 | 0.75324 | 10 | 11 | 13 | 5
GE.h | 80 | GE.h.080 | GE FX/i, LX/i | 16.33 | 6.7536 | 8.2208 | 16.33 | 1.4896 | 4.2 | 0.67615 | 0.3995 | 5 | 5 | 10 | 10
GE.h | 120 | GE.h.120 | GE FX/i, LX/i | 33.4 | 17.8752 | 19.4208 | 33.4 | 4.9952 | 10.1808 | 0.79656 | 0.63873 | 8 | 8 | 8 | 8
GE.h | 140 | GE.h.140 | GE FX/i, LX/i | 43.21 | 24.5056 | 25.872 | 43.21 | 7.2688 | 13.8656 | 0.8256 | 0.71579 | 9 | 11 | 9 | 5
GE.i | 80 | GE.i.080 | GE QX/i, LightSpeed, LightSpeed Plus | 13.2627 | 7.2366 | 7.5805 | 12.5004 | 1.6944 | 4.0493 | 0.79354 | 0.60061 | 7 | 6 | 8 | 7
GE.i | 100 | GE.i.100 | GE QX/i, LightSpeed, LightSpeed Plus | 22.1819 | 13.567 | 13.5457 | 20.9069 | 3.6611 | 7.6675 | 0.85627 | 0.76953 | 9 | 11 | 13 | 5
GE.i | 120 | GE.i.120 | GE QX/i, LightSpeed, LightSpeed Plus | 32.5438 | 21.0455 | 20.6477 | 27.3518 | 6.1934 | 12.0786 | 0.89202 | 1.0008 | 10 | 15 | 20 | 19
GE.i | 140 | GE.i.140 | GE QX/i, LightSpeed, LightSpeed Plus | 44.0987 | 29.5448 | 28.5304 | 41.5639 | 9.2297 | 17.1261 | 0.91312 | 0.96645 | 11 | 15 | 20 | 19
GE.j | 80 | GE.j.080 | GE HiSpeed ZX/i, NX/i | 8.3925 | 4.7702 | 4.9643 | 8.3925 | 1.1558 | 2.5592 | 0.82037 | 0.59637 | 9 | 6 | 9 | 7
GE.j | 120 | GE.j.120 | GE HiSpeed ZX/i, NX/i | 23.0227 | 14.7931 | 14.633 | 23.0227 | 4.411 | 7.7176 | 0.89097 | 0.80864 | 10 | 11 | 20 | 20
GE.j | 140 | GE.j.140 | GE HiSpeed ZX/i, NX/i | 32.4359 | 21.1224 | 20.1223 | 32.4359 | 6.4234 | 11.0459 | 0.88283 | 0.83508 | 10 | 11 | 13 | 20
GE.k | 80 | GE.k.080 | GE LightSpeed Ultra | 14.2375 | 7.7443 | 8.2333 | 10.5357 | 1.7838 | 4.484 | 0.79815 | 0.78597 | 8 | 11 | 8 | 20
GE.k | 100 | GE.k.100 | GE LightSpeed Ultra | 23.84 | 14.4977 | 14.7035 | 18.9098 | 3.8893 | 8.483 | 0.85951 | 0.93087 | 9 | 15 | 13 | 15
GE.k | 120 | GE.k.120 | GE LightSpeed Ultra | 34.9511 | 22.5476 | 22.3471 | 28.9809 | 6.9536 | 13.7924 | 0.89524 | 1.0709 | 10 | 19 | 20 | 12
GE.k | 140 | GE.k.140 | GE LightSpeed Ultra | 47.3621 | 31.6401 | 30.9318 | 40.52 | 9.8523 | 18.9306 | 0.91713 | 1.0769 | 11 | 19 | 20 | 12
GE.l | 80 | GE.l.080 | GE LightSpeed 16 | 13.0301 | 7.0895 | 7.5815 | 10.749 | 1.7166 | 3.8265 | 0.80108 | 0.70746 | 8 | 11 | 8 | 5
GE.l | 100 | GE.l.100 | GE LightSpeed 16 | 21.928 | 13.4396 | 13.6479 | 19.892 | 4.1803 | 8.1097 | 0.86631 | 0.921 | 10 | 15 | 13 | 15
GE.l | 120 | GE.l.120 | GE LightSpeed 16 | 32.1731 | 21.0608 | 20.8391 | 29.3377 | 7.1841 | 13.7708 | 0.90644 | 1.0845 | 10 | 19 | 20 | 12
GE.l | 140 | GE.l.140 | GE LightSpeed 16 | 43.8597 | 29.7086 | 28.9762 | 40.0141 | 10.7924 | 18.6042 | 0.92763 | 1.1707 | 11 | 20 | 20 | 16
GE.m | 80 | GE.m.080 | GE LightSpeed Pro 16 | 14.4946 | 7.4679 | 8.1014 | 10.0544 | 1.6604 | 4.1404 | 0.76897 | 0.7623 | 6 | 11 | 11 | 5
GE.m | 100 | GE.m.100 | GE LightSpeed Pro 16 | 23.2842 | 13.6676 | 13.9369 | 17.6068 | 3.5591 | 7.6728 | 0.83485 | 0.91008 | 9 | 15 | 9 | 15
GE.m | 120 | GE.m.120 | GE LightSpeed Pro 16 | 33.3009 | 20.9426 | 20.7581 | 26.647 | 6.018 | 11.908 | 0.87465 | 1.002 | 10 | 15 | 13 | 19
GE.m | 140 | GE.m.140 | GE LightSpeed Pro 16 | 44.4036 | 29.157 | 28.4293 | 37.0287 | 8.9496 | 16.6935 | 0.90139 | 1.0614 | 10 | 19 | 20 | 12
GE.n | 80 | GE.n.080 | GE LightSpeed RT | 13.1815 | 6.9294 | 7.9 | 9.1679 | 1.5793 | 4.2809 | 0.80643 | 0.8227 | 8 | 11 | 9 | 20
GE.n | 100 | GE.n.100 | GE LightSpeed RT | 21.113 | 12.552 | 13.5301 | 15.8807 | 3.3318 | 7.8162 | 0.87244 | 0.97321 | 10 | 15 | 13 | 19
GE.n | 120 | GE.n.120 | GE LightSpeed RT | 30.1575 | 19.2548 | 20.0744 | 23.8954 | 5.5974 | 12.0648 | 0.91322 | 1.0689 | 11 | 19 | 20 | 12
GE.n | 140 | GE.n.140 | GE LightSpeed RT | 42.1203 | 28.0273 | 28.6678 | 34.5841 | 8.6769 | 17.7091 | 0.93803 | 1.1331 | 11 | 19 | 17 | 16
GE.o | 80 | GE.o.080 | GE LightSpeed VCT | 14.8258 | 8.2762 | 9.9629 | 14.8258 | 1.9253 | 4.947 | 0.88031 | 0.5864 | 10 | 6 | 13 | 7
GE.o | 100 | GE.o.100 | GE LightSpeed VCT | 24.2325 | 15.3269 | 17.3511 | 24.2325 | 4.0405 | 8.7559 | 0.95092 | 0.73608 | 11 | 11 | 17 | 5
GE.o | 120 | GE.o.120 | GE LightSpeed VCT | 34.9667 | 23.7001 | 26.0597 | 34.9667 | 6.7256 | 13.2612 | 0.99591 | 0.83918 | 13 | 11 | 14 | 20
GE.o | 140 | GE.o.140 | GE LightSpeed VCT | 46.9416 | 33.118 | 35.5132 | 46.9416 | 9.9141 | 18.3306 | 1.0181 | 0.91389 | 15 | 15 | 19 | 15
GE.p | 80 | GE.p.080 | GE LightSpeed VCT (small hd, large bd) | 14.7783 | 7.7382 | 8.4484 | 10.4195 | 1.7354 | 4.4463 | 0.7832 | 0.7768 | 6 | 11 | 11 | 20
GE.p | 100 | GE.p.100 | GE LightSpeed VCT (small hd, large bd) | 24.1534 | 14.4214 | 14.9054 | 18.3898 | 3.7891 | 8.3492 | 0.85456 | 0.9356 | 9 | 15 | 13 | 15
GE.p | 120 | GE.p.120 | GE LightSpeed VCT (small hd, large bd) | 34.8586 | 22.3608 | 22.4808 | 27.8946 | 6.4536 | 13.0793 | 0.89796 | 1.0357 | 10 | 17 | 20 | 21
GE.p | 140 | GE.p.140 | GE LightSpeed VCT (small hd, large bd) | 46.8026 | 31.3231 | 30.9814 | 38.8141 | 9.6456 | 18.5321 | 0.92484 | 1.1026 | 11 | 19 | 20 | 16
PH.a | 120 | PH.a.120 | Philips 310, 350 (GE2, no Cu) | 32.8 | 18.7 | 21.2 |  |  |  | 0.86531 |  | 10 | #N/A | 13 | #N/A
PH.b | 120 | PH.b.120 | Philips 310, 350 (GE2, w. Cu) | 15.7648 | 11.1714 | 11.5887 |  |  |  | 1.0023 |  | 14 | #N/A | 14 | #N/A
PH.c | 120 | PH.c.120 | Philips 310, 350 (GE3, no Cu) |  |  |  | 21.6684 | 4.1999 | 10.2198 |  | 0.90296 | #N/A | 15 | #N/A | 15
PH.d | 120 | PH.d.120 | Philips 310, 350 (GE3, w. Cu) | [6 empty] | 1.1099 | 1.1439 | 20 | 20 | 16 | 16
PH.e | 80 | PH.e.080 | Philips AV, LX, SR7000 | 4.2565 | 2.5504 | 3.0077 | 8.72 | 1.38 | 3.52 | 0.92756 | 0.73247 | 11 | 11 | 20 | 5
PH.e | 100 | PH.e.100 | Philips AV, LX, SR7000 | 13.1906 | 8.8359 | 9.6427 | 13.1906 | 2.6065 | 5.7155 | 0.98069 | 0.89225 | 13 | 13 | 22 | 22
PH.e | 120 | PH.e.120 | Philips AV, LX, SR7000 | 19.2037 | 13.5515 | 14.7971 | 19.278 | 4.3352 | 9.0024 | 1.0294 | 1.0113 | 16 | 16 | 19 | 19
PH.e | 130 | PH.e.130 | Philips AV, LX, SR7000 | 22.5744 | 15.9989 | 17.5846 | 22.781 | 5.2586 | 11.0876 | 1.0377 | 1.0451 | 17 | 18 | 23 | 23
PH.e | 140 | PH.e.140 | Philips AV, LX, SR7000 | 25.9589 | 18.9553 | 20.3161 | 25.9589 | 6.2095 | 12.413 | 1.0508 | 1.0697 | 18 | 19 | 21 | 12
PH.f | 120 | PH.f.120 | Philips CX, CX/S | 20.5061 | 14.1558 | 15.4081 | 19.2391 | 3.9998 | 7.4304 | 1.0068 | 0.89934 | 15 | 14 | 18 | 18
PH.g | 120 | PH.g.120 | Philips SR4000 | 18.2037 | 12.5065 | 13.5097 | 18.2037 | 3.7578 | 7.8362 | 0.99777 | 0.92209 | 13 | 15 | 14 | 15
PH.h | 120 | PH.h.120 | Philips SR 5000 | 18.0186 | 11.8446 | 13.0028 | 18.0186 | 3.5685 | 7.6778 | 0.9672 | 0.88927 | 12 | 12 | 17 | 22
PH.h | 130 | PH.h.130 | Philips SR 5000 | 20.8472 | 14.0534 | 15.1451 | 20.8472 | 4.3558 | 10.1423 | 0.97905 | 0.96655 | 12 | 15 | 22 | 19
PH.i | 120 | PH.i.120 | Philips M / EG | 58.4 | 34.8 | 45.2 | 58.4 | 11.3 | 39.1 | 0.98019 | 1.027 | 12 | 17 | 22 | 21
PH.i | 130 | PH.i.130 | Philips M / EG | 67.6 | 41.6 | 53.8 | 67.6 | 13.9 | 47.9 | 1.007 | 1.0952 | 15 | 19 | 18 | 12
PH.j | 100 | PH.j.100 | Philips TX | [6 empty] | 0.91 | 0.86 | 11 | 12 | 20 | 20
PH.j | 120 | PH.j.120 | Philips TX | 7.8 | 5.2 | 5.5 |  |  |  | 0.95834 |  | 12 |  | 17 | 17
PH.j | 130 | PH.j.130 | Philips TX | [6 empty] | 1.0465 | 1.0266 | 18 | 17 | 21 | 21
PH.k | 120 | PH.k.120 | Philips CT Secura | 19.5052 | 13.4422 | 14.6592 | 19.5052 | 4.4972 | 9.1354 | 1.0063 | 1.0326 | 14 | 17 | 18 | 21
PH.k | 140 | PH.k.140 | Philips CT Secura | 27.4497 | 19.2981 | 20.801 | 27.4497 | 6.7156 | 13.1378 | 1.0179 | 1.0895 | 15 | 19 | 18 | 12
PH.l | 90 | PH.l.090 | Philips Mx8000 | 9.1314 | 6.4111 | 7.3315 | 9.1314 | 1.7875 | 4.1673 | 1.0538 | 0.90019 | 18 | 14 | 21 | 18
PH.l | 120 | PH.l.120 | Philips Mx8000 | 19.8994 | 14.7243 | 16.0822 | 19.8994 | 4.6533 | 9.4223 | 1.076 | 1.0476 | 18 | 18 | 12 | 23
PH.l | 140 | PH.l.140 | Philips Mx8000 | 28.6207 |  |  | 28.6207 | 7.022 | 13.8987 |  | 1.0965 |  | 19 |  | 12
PH.m | 120 | PH.m.120 | Philips AcQSim | 27.712 | 15.3231 | 18.3185 | 27.712 | 4.5522 | 11.7713 | 0.86898 | 0.76737 | 10 | 11 | 13 | 5
PH.m | 130 | PH.m.130 | Philips AcQSim | 31.2523 | 17.9856 | 21.0711 | 31.2523 | 5.5205 | 13.6588 | 0.89029 | 0.81948 | 10 | 11 | 20 | 20
PH.n | 90 | PH.n.090 | Mx8000 IDT / Brilliance 16 (& Power) | 9.1411 | 6.6375 | 7.3508 | 9.145 | 1.986 | 4.2646 | 1.0662 | 0.98328 | 18 | 15 | 12 | 19
PH.n | 120 | PH.n.120 | Mx8000 IDT / Brilliance 16 (& Power) | 19.4504 | 14.6855 | 15.992 | 19.4586 | 4.8679 | 9.4181 | 1.0944 | 1.1127 | 19 | 19 | 12 | 16
PH.n | 140 | PH.n.140 | Mx8000 IDT / Brilliance 16 (& Power) | 28.0698 | 21.2993 | 23.2671 | 28.0817 | 7.308 | 13.3776 | 1.1016 | 1.144 | 19 | 20 | 16 | 16
PH.o | 120 | PH.o.120 | Philips Aura | 39.5747 | 21.7745 | 25.5 | 39.5747 | 7.1779 | 14.3559 | 0.85427 | 0.78944 | 9 | 11 | 13 | 20
PH.p | 90 | PH.p.090 | Philips Big Bore | 6.9024 | 4.2557 | 4.978 | 6.9814 | 1.3235 | 3.3001 | 0.94753 | 0.88842 | 11 | 12 | 17 | 22
PH.p | 120 | PH.p.120 | Philips Big Bore | 14.8685 | 9.8575 | 11.0907 | 14.9159 | 3.3611 | 7.2844 | 0.98942 | 1.0265 | 13 | 16 | 22 | 21
PH.p | 140 | PH.p.140 | Philips Big Bore | 21.3203 | 14.548 | 15.93 | 21.2534 | 5.1169 | 10.625 | 0.9996 | 1.0891 | 13 | 19 | 14 | 12
PH.q | 80 | PH.q.080 | Philips Brilliance 64 or 40 | 6.982 | 4.591 | 5.397 | 6.982 | 1.2911 | 3.0064 | 1.0086 | 0.84506 | 15 | 11 | 18 | 20
PH.q | 120 | PH.q.120 | Philips Brilliance 64 or 40 | 21.4527 | 15.6666 | 16.8468 | 21.4527 | 5.0523 | 9.993 | 1.053 | 1.0487 | 18 | 18 | 21 | 23
PH.q | 140 | PH.q.140 | Philips Brilliance 64 or 40 | 30.9822 | 22.9059 | 24.7677 | 30.9822 | 7.8746 | 14.8995 | 1.0686 | 1.1251 | 18 | 19 | 12 | 16
PI.a | 80 | PI.a.080 | Picker 1200SX |  |  |  | 16.2587 | 1.3167 | 5.2826 |  | 0.40567 |  | 5 |  | 10
PI.a | 120 | PI.a.120 | Picker 1200SX | 30.3355 | 12.1255 | 11.2189 | 30.0328 | 4.2507 | 11.1803 | 0.56211 | 0.65266 | 3 | 8 | 4 | 11
PI.a | 130 | PI.a.130 | Picker 1200SX | 31.7871 | 12.2689 | 12.5962 | 32.5937 | 4.8819 | 12.5962 | 0.57687 | 0.6912 | 4 | 11 | 4 | 5
PI.a | 140 | PI.a.140 | Picker 1200SX | 37.9707 | 16.9155 | 14.2525 | 38.778 | 6.4201 | 15.4554 | 0.58824 | 0.75542 | 4 | 11 | 4 | 5
PI.b | 120 | PI.b.120 | Picker PQ Series | 30.1015 | 13.9917 | 13.2855 | 30.0475 | 4.8459 | 11.825 | 0.6505 | 0.73687 | 5 | 11 | 6 | 5
PI.b | 130 | PI.b.130 | Picker PQ Series | 36.85 | 17.4663 | 16.1639 | 36.7833 | 6.0143 | 15.5127 | 0.65266 | 0.76271 | 5 | 11 | 6 | 5
PI.b | 140 | PI.b.140 | Picker PQ Series | 38.1523 | 19.1239 | 17.6755 | 37.9783 | 6.9769 | 16.7795 | 0.68541 | 0.84782 | 6 | 11 | 10 | 20
PI.c | 80 | PI.c.080 | Picker UltraZ | 16.5 | 5.55 | 6.18 | 17.655 | 1.16 | 5.21 | 0.53589 | 0.33204 | 3 | 5 | 4 | 10
PI.c | 100 | PI.c.100 | Picker UltraZ | 23.66 | 9.65 | 10.33 | 25.3162 | 3.02 | 8.39 | 0.61969 | 0.54708 | 4 | 6 | 6 | 10
PI.c | 120 | PI.c.120 | Picker UltraZ | 31.44 | 14.34 | 13.84 | 33.6408 | 4.91 | 11.94 | 0.64545 | 0.65752 | 5 | 8 | 6 | 11
PI.c | 130 | PI.c.130 | Picker UltraZ | 35.58 | 16.87 | 15.98 | 38.0706 | 5.96 | 14.24 | 0.66117 | 0.7076 | 5 | 11 | 10 | 5
PI.c | 140 | PI.c.140 | Picker UltraZ | 39.69 | 19.5 | 18.33 | 42.4683 | 7.07 | 16.28 | 0.67952 | 0.74907 | 6 | 11 | 10 | 5
PI.d | 90 | PI.d.090 | Philips / Marconi Mx8000 | 9.1314 | 6.4111 | 7.3315 | 9.1314 | 1.7875 | 4.1673 | 1.0538 | 0.90019 | 18 | 14 | 21 | 18
PI.d | 120 | PI.d.120 | Philips / Marconi Mx8000 | 19.8994 | 14.7243 | 16.0822 | 19.8994 | 4.6533 | 9.4223 | 1.076 | 1.0476 | 18 | 18 | 12 | 23
PI.d | 140 | PI.d.140 | Philips / Marconi Mx8000 | 28.6207 |  |  | 28.6207 | 7.022 | 13.8987 |  | 1.0965 |  | 19 |  | 12
PI.e | 120 | PI.e.120 | Philips / Marconi AcQSim | 27.712 | 15.3231 | 18.3185 | 27.712 | 4.5522 | 11.7713 | 0.86898 | 0.76737 | 10 | 11 | 13 | 5
PI.e | 130 | PI.e.130 | Philips / Marconi AcQSim | 31.2523 | 17.9856 | 21.0711 | 31.2523 | 5.5205 | 13.6588 | 0.89029 | 0.81948 | 10 | 11 | 20 | 20
SH.a | 80 | SH.a.080 | Shimadzu SCT | 5.55 | 3.6346 | 4.3654 | 5.4219 | 1.0188 | 3.2656 | 1.0183 | 0.9644 | 15 | 15 | 19 | 19
SH.a | 120 | SH.a.120 | Shimadzu SCT | 15.5962 | 12.0064 | 13.4231 | 15.5937 | 4.0469 | 10.0664 | 1.1324 | 1.2485 | 21 | 20 | 3 | 1
SH.a | 130 | SH.a.130 | Shimadzu SCT | 17.4 | 13.5 | 14.9231 | 17.3125 | 4.625 | 11.4531 | 1.1328 | 1.2859 | 21 | 20 | 3 | 1
SI.a | 125 | SI.a.125 | Siemens CR, CR512, DRH | 12.4489 | 9.2598 | 10.9449 | 13.2178 | 3.4919 | 9.5863 | 1.1349 | 1.3156 | 22 | 22 | 3 | 3
SI.b | 125 | SI.b.125 | Siemens Somatom 2, DR1/2/3 | 8.8771 | 6.4969 | 7.6367 | 8.8771 | 2.332 | 6.4949 | 1.1141 | 1.3143 | 21 | 21 | 1 | 1
SI.c | 125 | SI.c.125 | Siemens DRG, DRG1 | [6 empty] | 1.2186 | 1.4272 | 23 | 23 | 2 | 2
SI.d | 80 | SI.d.080 | Siemens Somatom Plus 4 Series | 5.6288 | 3.7341 | 4.2929 | 5.6288 | 1.1456 | 2.9456 | 1.0031 | 0.97044 | 14 | 15 | 14 | 19
SI.d | 120 | SI.d.120 | Siemens Somatom Plus 4 Series | 17.7128 | 13.5549 | 15.0717 | 17.8675 | 4.4362 | 9.6412 | 1.1223 | 1.1412 | 21 | 19 | 1 | 16
SI.d | 140 | SI.d.140 | Siemens Somatom Plus 4 Series | 25.0829 | 19.455 | 21.2542 | 25.1829 | 6.5784 | 13.8573 | 1.1244 | 1.1943 | 21 | 20 | 1 | 16
SI.e | 110 | SI.e.110 | Siemens Somatom AR-C, AR.SP, AR-T | 24.4564 | 16.2491 | 18.1125 | 24.4564 | 4.8884 | 10.8819 | 0.98581 | 0.90778 | 13 | 15 | 22 | 15
SI.e | 130 | SI.e.130 | Siemens Somatom AR-C, AR.SP, AR-T | 35.7449 | 25.1984 | 27.7336 | 35.7387 | 7.8128 | 16.4426 | 1.0334 | 0.98447 | 16 | 15 | 19 | 19
SI.f | 130 | SI.f.130 | Siemens AR.HP | 33.5 | 26.18 | 27.6 | 33.5 | 8.88 | 16.4 | 1.1083 | 1.1697 | 19 | 20 | 16 | 16
SI.g | 120 | SI.g.120 | Siemens Plus, DXP, Plus-S | 12.2372 | 9.4 | 10.84 | 12.2372 | 3.2 | 6.985 | 1.1518 | 1.2083 | 22 | 20 | 3 | 16
SI.g | 137 | SI.g.137 | Siemens Plus, DXP, Plus-S | 15.55 | 12.2247 | 13.48 | 16.17 | 4.258 | 9.04 | 1.1451 | 1.2074 | 22 | 20 | 3 | 16
SI.h | 133 | SI.h.133 | Siemens Hi Q | 17.6369 | 12.9264 | 14.4601 | 17.6369 | 4.0399 | 10.2623 | 1.082 | 1.099 | 18 | 19 | 12 | 12
SI.i | 110 | SI.i.110 | Siemens Balance, Emotion | 23.6956 | 15.3064 | 17.262 | 23.6956 | 4.6545 | 10.2843 | 0.96733 | 0.88849 | 12 | 12 | 17 | 22
SI.i | 130 | SI.i.130 | Siemens Balance, Emotion | 33.1798 | 22.4654 | 24.9491 | 33.1798 | 7.2824 | 15.2426 | 1.0009 | 0.98717 | 13 | 15 | 14 | 19
SI.j | 80 | SI.j.080 | Siemens Volume Zoom, Access | 10.7256 | 6.336 | 8.2457 | 6.265 | 1.236 | 3.3407 | 0.97358 | 0.95435 | 12 | 15 | 22 | 15
SI.j | 120 | SI.j.120 | Siemens Volume Zoom, Access | 24.9904 | 17.595 | 20.8617 | 17.4679 | 4.557 | 9.6861 | 1.0804 | 1.1957 | 18 | 20 | 12 | 16
SI.j | 140 | SI.j.140 | Siemens Volume Zoom, Access | 34.097 | 24.6763 | 28.6417 | 24.418 | 6.8379 | 13.974 | 1.0939 | 1.2756 | 19 | 20 | 12 | 1
SI.k | 80 | SI.k.080 | Siemens Emotion Duo | 9.3907 | 5.4673 | 6.3531 | 9.3907 | 1.3236 | 3.2116 | 0.89532 | 0.63141 | 10 | 7 | 20 | 8
SI.k | 110 | SI.k.110 | Siemens Emotion Duo | 20.9866 | 13.767 | 14.8968 | 20.9866 | 3.927 | 7.8505 | 0.95706 | 0.81718 | 11 | 11 | 17 | 20
SI.k | 130 | SI.k.130 | Siemens Emotion Duo | 30.2348 | 20.6981 | 21.9071 | 30.2348 | 6.2318 | 11.8947 | 0.98247 | 0.89751 | 13 | 13 | 22 | 18
SI.l | 80 | SI.l.080 | Siemens Sensation 4 | 12.1751 | 7.2257 | 8.9123 | 6.9429 | 1.4246 | 3.4321 | 0.94529 | 0.95805 | 11 | 15 | 17 | 19
SI.l | 120 | SI.l.120 | Siemens Sensation 4 | 27.227 | 18.9815 | 21.4215 | 18.8814 | 5.061 | 10.0431 | 1.0385 | 1.2071 | 17 | 20 | 23 | 16
SI.l | 140 | SI.l.140 | Siemens Sensation 4 |  |  |  | 26.4025 | 7.455 | 14.1814 |  | 1.2617 |  | 20 |  | 1
SI.m | 80 | SI.m.080 | Siemens Sensation 16 | 9.063 | 5.521 | 6.6988 | 5.2428 | 1.1171 | 2.6136 | 0.95847 | 0.98899 | 12 | 15 | 17 | 19
SI.m | 100 | SI.m.100 | Siemens Sensation 16 | 16.0831 | 10.846 | 12.5291 | 10.5944 | 2.4691 | 5.22 | 1.0214 | 1.0569 | 15 | 19 | 19 | 12
SI.m | 120 | SI.m.120 | Siemens Sensation 16 | 21.752 | 15.3636 | 17.1829 | 15.2625 | 4.1664 | 8.1595 | 1.0454 | 1.2265 | 17 | 20 | 21 | 16
SI.m | 140 | SI.m.140 | Siemens Sensation 16 |  |  |  | 21.63 | 6.1902 | 11.6165 |  | 1.2754 |  | 20 |  | 1
SI.n | 80 | SI.n.080 | Siemens Sensation 16 Straton | 9.5936 | 5.3646 | 7.4392 | 5.6791 | 1.1135 | 2.6051 | 0.96397 | 0.90281 | 12 | 14 | 17 | 15
SI.n | 100 | SI.n.100 | Siemens Sensation 16 Straton | 16.5563 | 10.1572 | 13.3406 | 10.9159 | 2.5222 | 5.0311 | 1.0141 | 1.0296 | 15 | 17 | 18 | 21
SI.n | 120 | SI.n.120 | Siemens Sensation 16 Straton | 24.2558 | 15.6176 | 17.6682 | 17.1653 | 4.3118 | 8.3771 | 0.96628 | 1.1189 | 12 | 19 | 17 | 16
SI.n | 140 | SI.n.140 | Siemens Sensation 16 Straton |  |  |  | 24.7924 | 6.5533 | 12.1659 |  | 1.1677 |  | 20 |  | 16
SI.o | 80 | SI.o.080 | Siemens Emotion 6 | 9.1858 | 5.34 | 5.8208 | 9.1858 | 1.4639 | 3.1134 | 0.86042 | 0.69546 | 9 | 11 | 13 | 5
SI.o | 110 | SI.o.110 | Siemens Emotion 6 | 19.4222 | 12.81 | 13.4581 | 19.4222 | 4.0024 | 7.5114 | 0.94516 | 0.89314 | 11 | 13 | 17 | 22
SI.o | 130 | SI.o.130 | Siemens Emotion 6 | 26.5911 | 18.425 | 19.0816 | 26.5911 | 6.1153 | 10.9403 | 0.9808 | 0.99443 | 13 | 15 | 22 | 19
SI.p | 80 | SI.p.080 | Siemens Sensation 10 | 8.8945 | 5.3784 | 6.625 | 5.275 | 1.1074 | 2.5909 | 0.96092 | 0.97306 | 12 | 15 | 17 | 19
SI.p | 100 | SI.p.100 | Siemens Sensation 10 | 15.4041 | 10.2348 | 11.9397 | 10.1787 | 2.3372 | 4.9412 | 1.0136 | 1.04 | 15 | 17 | 18 | 23
SI.p | 120 | SI.p.120 | Siemens Sensation 10 | 20.2916 | 14.1206 | 15.9485 | 14.4026 | 3.8737 | 7.5862 | 1.0372 | 1.2071 | 16 | 20 | 23 | 16
SI.p | 140 | SI.p.140 | Siemens Sensation 10 |  |  |  | 20.8906 | 5.8904 | 11.0539 |  | 1.2553 |  | 20 |  | 1
SI.q | 80 | SI.q.080 | Siemens Sensation Open | 8.4486 | 5.0464 | 5.5852 | 8.4761 | 1.3098 | 2.9057 | 0.89005 | 0.68058 | 10 | 9 | 20 | 9
SI.q | 100 | SI.q.100 | Siemens Sensation Open | 14.9407 | 9.8375 | 10.515 | 15.0293 | 2.8505 | 5.7332 | 0.95336 | 0.83099 | 11 | 11 | 17 | 20
SI.q | 120 | SI.q.120 | Siemens Sensation Open | 22.4558 | 15.535 | 16.4005 | 22.6494 | 4.8331 | 9.1898 | 0.99054 | 0.93138 | 13 | 15 | 22 | 15
SI.q | 140 | SI.q.140 | Siemens Sensation Open |  |  |  | 32.2413 | 7.4148 | 13.6156 |  | 1.0013 |  | 15 |  | 19
SI.r | 80 | SI.r.080 | Siemens Sensation 64 | 4.9457 | 3.2553 | 3.4594 | 4.9473 | 0.89006 | 1.8034 | 0.94979 | 0.7853 | 11 | 11 | 17 | 20
SI.r | 100 | SI.r.100 | Siemens Sensation 64 | 9.9314 | 7.0129 | 7.2921 | 9.9432 | 2.1315 | 3.9886 | 1.0005 | 0.93196 | 13 | 15 | 14 | 15
SI.r | 120 | SI.r.120 | Siemens Sensation 64 | 16.1277 | 11.7008 | 12.0852 | 16.145 | 3.7965 | 6.8243 | 1.0218 | 1.0201 | 15 | 16 | 19 | 21
SI.r | 140 | SI.r.140 | Siemens Sensation 64 |  |  |  | 24.5712 | 6.0992 | 11.024 |  | 1.0834 |  | 19 |  | 12
SI.s | 80 | SI.s.080 | Siemens Definition AS | 5.7316 | 3.34 | 3.61 | 5.7209 | 0.87 | 1.98 | 0.858 | 0.67387 | 9 | 9 | 13 | 11
SI.s | 100 | SI.s.100 | Siemens Definition AS | 10.4589 | 6.69 | 7.05 | 10.4375 | 1.94 | 3.92 | 0.92055 | 0.81365 | 11 | 11 | 20 | 20
SI.s | 120 | SI.s.120 | Siemens Definition AS | 16.9485 | 11.32 | 11.84 | 16.9217 | 3.52 | 6.62 | 0.95367 | 0.90293 | 11 | 14 | 17 | 15
SI.s | 140 | SI.s.140 | Siemens Definition AS | 24.33 | 16.6 | 17.28 | 24.346 | 5.45 | 9.74 | 0.96985 | 0.96531 | 12 | 15 | 22 | 19
TO.a | 120 | TO.a.120 | Toshiba TCT 600 | 22.2619 | 16.0795 | 15.767 | 18.26 | 3.8929 | 6.9702 | 0.98721 | 0.91547 | 13 | 15 | 22 | 15
TO.b | 120 | TO.b.120 | Toshiba Xspeed II | 30.69 | 21.38 | 21.22 | 11.81 | 2.9101 | 5.5197 | 0.96153 | 1.0887 | 12 | 19 | 17 | 12
TO.c | 120 | TO.c.120 | Toshiba Xpress GX (Pre '98) | 33.61 | 17.75 | 18.67 |  |  |  | 0.77231 |  | 6 |  | 11
TO.d | 120 | TO.d.120 | Toshiba Xvision / EX | 25.5956 | 16.9903 | 15.7628 | 21.9531 | 4.8355 | 7.405 | 0.88515 | 0.91273 | 10 | 15 | 13 | 15
TO.e | 120 | TO.e.120 | Toshiba Xpress HS1 | 30.57 | 19.66 | 20.82 | 22.01 | 5.28 | 8.2 | 0.92782 | 1.0054 | 11 | 15 | 20 | 19
TO.f | 120 | TO.f.120 | Toshiba Xpress HS | 17.38 | 12.9122 | 12.9122 | 21.97 | 4.7332 | 7.281 | 1.0249 | 0.89169 | 15 | 12 | 19 | 22
TO.g | 120 | TO.g.120 | Toshiba Xpress GX (Post '98), Asteion | 19.78 | 13.04 | 13.72 | 20.9541 | 4.0275 | 7.055 | 0.94558 | 0.81176 | 11 | 11 | 17 | 20
TO.g | 130 | TO.g.130 | Toshiba Xpress GX (Post '98), Asteion | 23.09 | 15.24 | 15.64 | 23.5688 | 4.8733 | 8.325 | 0.93285 | 0.87442 | 11 | 12 | 20 | 20
TO.h | 80 | TO.h.080 | Toshiba Aquilion Multi / 4 | 14.6509 | 7.597 | 8.9345 | 20.3364 | 2.1597 | 5.633 | 0.81149 | 0.46572 | 8 | 5 | 9 | 10
TO.h | 100 | TO.h.100 | Toshiba Aquilion Multi / 4 | 23.4627 | 13.8441 | 15.4932 | 30.7922 | 4.2896 | 9.7341 | 0.886 | 0.60915 | 10 | 7 | 13 | 7
TO.h | 120 | TO.h.120 | Toshiba Aquilion Multi / 4 | 33.8449 | 21.4703 | 23.3086 | 42.4807 | 7.0726 | 14.799 | 0.92982 | 0.72698 | 11 | 11 | 20 | 5
TO.h | 135 | TO.h.135 | Toshiba Aquilion Multi / 4 | 43.5553 | 28.6935 | 30.1461 | 52.0748 | 9.8068 | 19.6934 | 0.94415 | 0.82409 | 11 | 11 | 17 | 20
TO.i | 120 | TO.i.120 | Toshiba Auklet | 36.9 | 18 | 18.5 | 36.9 | 4.4 | 7.5 | 0.70966 | 0.4658 | 6 | 5 | 10 | 10
TO.j | 80 | TO.j.080 | Toshiba Asteion Multi (older tube) | 26.323 | 11.6993 | 14.1777 | 26.2251 | 2.7119 | 7.3128 | 0.71909 | 0.45689 | 6 | 5 | 10 | 10
TO.j | 100 | TO.j.100 | Toshiba Asteion Multi (older tube) | 40.3721 | 21.155 | 24.2903 | 40.2219 | 5.6717 | 14.5088 | 0.80751 | 0.64347 | 8 | 8 | 9 | 8
TO.j | 120 | TO.j.120 | Toshiba Asteion Multi (older tube) | 56.0056 | 32.4629 | 36.1689 | 55.7973 | 9.4996 | 19.9136 | 0.86938 | 0.74586 | 10 | 11 | 13 | 5
TO.j | 135 | TO.j.135 | Toshiba Asteion Multi (older tube) | 68.3595 | 41.7329 | 45.6298 | 68.1052 | 12.8473 | 29.0633 | 0.90145 | 0.85596 | 10 | 11 | 20 | 20
TO.k | 80 | TO.k.080 | Toshiba Asteion Multi (CXB-400C tube) | 21.1168 | 9.3854 | 11.3736 | 21.0382 | 2.1755 | 5.8665 | 0.71909 | 0.45689 | 6 | 5 | 10 | 10
TO.k | 100 | TO.k.100 | Toshiba Asteion Multi (CXB-400C tube) | 31.9739 | 16.7543 | 19.2374 | 31.8549 | 4.4919 | 11.4907 | 0.80751 | 0.64347 | 8 | 8 | 9 | 8
TO.k | 120 | TO.k.120 | Toshiba Asteion Multi (CXB-400C tube) | 44.1109 | 25.5683 | 28.4872 | 43.9468 | 7.482 | 15.6843 | 0.86938 | 0.74586 | 10 | 11 | 13 | 5
TO.k | 135 | TO.k.135 | Toshiba Asteion Multi (CXB-400C tube) | 54.0731 | 33.0112 | 36.0936 | 53.872 | 10.1624 | 22.9894 | 0.90145 | 0.85596 | 10 | 11 | 20 | 20
TO.l | 120 | TO.l.120 | Toshiba Asteion Dual | 46.5472 | 26.0311 | 30.6169 | 45.9693 | 8.0539 | 18.4066 | 0.86934 | 0.79114 | 10 | 11 | 13 | 20
TO.l | 135 | TO.l.135 | Toshiba Asteion Dual | 56.901 | 33.5791 | 37.1193 | 56.1946 | 10.9448 | 22.1868 | 0.87962 | 0.85772 | 10 | 11 | 13 | 20
TO.m | 80 | TO.m.080 | Toshiba Aquilion 16 | 15.9132 | 8.25 | 10.0645 | 21.9265 | 2.3623 | 6.637 | 0.82965 | 0.48749 | 9 | 5 | 9 | 10
TO.m | 100 | TO.m.100 | Toshiba Aquilion 16 | 25.275 | 15.6654 | 17.3193 | 33.1999 | 4.7883 | 11.6741 | 0.92013 | 0.64925 | 11 | 8 | 20 | 11
TO.m | 120 | TO.m.120 | Toshiba Aquilion 16 | 36.1535 | 24.5332 | 26.6036 | 45.8023 | 8.029 | 17.4089 | 0.98871 | 0.77862 | 13 | 11 | 22 | 20
TO.m | 135 | TO.m.135 | Toshiba Aquilion 16 | 45.4255 | 32.4563 | 34.9504 | 56.1466 | 11.0254 | 23.0583 | 1.0327 | 0.8735 | 16 | 12 | 19 | 20
Note on D51: This entry for the LightSpeed VCT covers all scan modes for which the 'medium' beam shaping filter is employed (i.e. except where small head, large body or small cardiac filters are used).
Note on D55: This entry covers the LightSpeed VCT for scan modes which use the 'small' beam shaping filter for head scans and the 'large' filter for body scans. Cardiac scans using the 'small' filter are not currently covered by these data.
Note on D82: CTDI for 8mm collimation (10 mm not available)
Note on D83: CTDI for 8mm collimation (10 mm not available)
Note on D84: CTDI at 12mm collimation
Note on D85: CTDI at 12mm collimation
Note on D86: CTDI at 12mm collimation
Note on D109: CTDI for 8mm collimation (10 mm not available)
Note on D110: CTDI for 8mm collimation (10 mm not available)
Note on D172: CTDI values given for 8mm collimation
Note on D173: CTDI values given for 8mm collimation
Note on D174: CTDI values given for 8mm collimation
Note on D175: CTDI values given for 8mm collimation
Note on D177: CTDI values for 8mm collimation
Note on D178: CTDI values for 8mm collimation
Note on D179: CTDI values for 8mm collimation
Note on D180: CTDI values for 8mm collimation
Note on D181: CTDI values for 8mm collimation
Note on D182: CTDI values for 8mm collimation
Note on D183: CTDI values for 8mm collimation
Note on D184: CTDI values for 8mm collimation
Note on D187: CTDI values for 8mm collimation
Note on D188: CTDI values for 8mm collimation
Note on D189: CTDI values for 8mm collimation
Note on D190: CTDI values for 8mm collimation
Note on M1: Intermediate step used for scanner matching

## MatchData
ImFHd | Scanner | MC Data Set | ImFBd | Scanner | MC Data Set
0 | Picker 1200SX Body | 5 | 0 | Philips TX, 120kV | 17
0 | GE 8800, 9000 I/II/HP Body | 7 | 0 | Picker 1200SX Head | 4
0 | Philips 305N, 310, 350 (GE3, No Cu) | 15 | 0 | GE 8800, 9000 I/II/HP Head | 6
0.57687 | Picker 1200SX Head | 4 | 0 | Philips 305N, 310, 350 (GE2, No Cu) | 13
0.64065 | GE 8800, 9000 I/II/HP Head | 6 | 0 | Philips 305N, 310, 350 (GE2, with Cu) | 14
0.67706 | GE Max | 10 | 0.51965 | GE Max | 10
0.78796 | GE Pace | 11 | 0.60617 | GE 8800, 9000 I/II/HP Body | 7
0.79429 | GE 9800, 9800 Quick, 120kV | 8 | 0.63416 | GE 9800, 9800 Quick, 120kV | 8
0.81554 | GE 9800, 9800 Quick, 140kV | 9 | 0.66367 | GE Pace | 11
0.86531 | Philips 305N, 310, 350 (GE2, No Cu) | 13 | 0.69018 | GE 9800, 9800 Quick, 140kV | 9
0.91 | Philips TX, 100kV | 20 | 0.6912 | Picker 1200SX Body | 5
0.95834 | Philips TX, 120kV | 17 | 0.86 | Philips TX, 100kV | 20
0.98069 | Philips LX, 100kV | 22 | 0.89225 | Philips LX, 100kV | 22
1.0023 | Philips 305N, 310, 350 (GE2, with Cu) | 14 | 0.89934 | Philips CX, CX/S | 18
1.0068 | Philips CX, CX/S | 18 | 0.90296 | Philips 305N, 310, 350 (GE3, No Cu) | 15
1.0294 | Philips LX, 120kV | 19 | 1.0113 | Philips LX, 120kV | 19
1.0377 | Philips LX, 130kV | 23 | 1.0266 | Philips TX, 130kV | 21
1.0465 | Philips TX, 130kV | 21 | 1.0451 | Philips LX, 130kV | 23
1.0827 | CGR CE 10000, 12000 | 12 | 1.0559 | CGR CE 10000, 12000 | 12
1.1099 | Philips 305N, 310, 350 (GE3, with Cu) | 16 | 1.1439 | Philips 305N, 310, 350 (GE3, with Cu) | 16
1.1141 | Siemens Somatom 2, DR1/2/3 | 1 | 1.3143 | Siemens Somatom 2, DR1/2/3 | 1
1.1349 | Siemens Somatom DRH,CR,CR512 | 3 | 1.3156 | Siemens Somatom DRH,CR,CR512 | 3
1.2186 | Siemens Somatom DRG, DRG1 | 2 | 1.4272 | Siemens Somatom DRG, DRG1 | 2
Figures in Red are interpolated ImPACT Factors

## Collimation
Scanner Group | Collimation setting | Sub-group | Collimation (mm) | Rel. CTDI
EL.a | 1 | EL.a.1 | 10 | 1
EL.a | 2 | EL.a.2 | 5 | 1.0598
EL.a | 3 | EL.a.3 | 2.5 | 1.0684
EL.a | 4 | EL.a.4 | 1.5 | 1.265
EL.b | 1 | EL.b.1 | 20 | 0.94
EL.b | 2 | EL.b.2 | 16 | 0.98
EL.b | 3 | EL.b.3 | 10 | 1
EL.b | 4 | EL.b.4 | 5 | 1.1
EL.b | 5 | EL.b.5 | 2 | 1.62
EL.b | 6 | EL.b.6 | 1 | 1.81
GE.c | 1 | GE.c.1 | 10 | 1
GE.c | 2 | GE.c.2 | 7 | 1.0059
GE.c | 3 | GE.c.3 | 5 | 1.0178
GE.c | 4 | GE.c.4 | 3 | 1.0355
GE.c | 5 | GE.c.5 | 1 | 1.0473
GE.f | 1 | GE.f.1 | 10 | 1
GE.f | 2 | GE.f.2 | 5 | 0.99
GE.f | 3 | GE.f.3 | 3 | 1.01
GE.f | 4 | GE.f.4 | 1 | 1.26
GE.g | 1 | GE.g.1 | 10 | 1
GE.g | 2 | GE.g.2 | 5 | 1.0385
GE.g | 3 | GE.g.3 | 3 | 1.0275
GE.g | 4 | GE.g.4 | 1 | 1.1209
GE.h | 1 | GE.h.1 | 10 | 1
GE.h | 2 | GE.h.2 | 7 | 0.99
GE.h | 3 | GE.h.3 | 5 | 0.98
GE.h | 4 | GE.h.4 | 3 | 0.98
GE.h | 5 | GE.h.5 | 2 | 0.98
GE.h | 6 | GE.h.6 | 1 | 1.12
GE.i | 1 | GE.i.1 | 20 | 0.89172
GE.i | 2 | GE.i.2 | 15 | 0.9869
GE.i | 3 | GE.i.3 | 10 | 1
GE.i | 4 | GE.i.4 | 5 | 1.2603
GE.i | 5 | GE.i.5 | 1x1.25 | 2.1175
GE.i | 6 | GE.i.6 | 2x.63 | 1.308
GE.j | 1 | GE.j.1 | 20 | 0.9846
GE.j | 2 | GE.j.2 | 14 | 0.99513
GE.j | 3 | GE.j.3 | 10 | 1
GE.j | 4 | GE.j.4 | 7 | 0.98524
GE.j | 5 | GE.j.5 | 6 | 1.0216
GE.j | 6 | GE.j.6 | 5 | 0.98921
GE.j | 7 | GE.j.7 | 4 | 1.0231
GE.j | 8 | GE.j.8 | 3 | 0.9776
GE.j | 9 | GE.j.9 | 2 (NX/i) | 1.0686
GE.j | A | GE.j.A | 2 (ZX/i) | 0.97805
GE.j | B | GE.j.B | 1 (NX/i) | 0.79408
GE.j | C | GE.j.C | 1 (ZX/i) | 1.0964
GE.k | 1 | GE.k.1 | 20 | 0.87112
GE.k | 2 | GE.k.2 | 15 | 1.0147
GE.k | 3 | GE.k.3 | 10 | 1
GE.k | 4 | GE.k.4 | 5 | 1.2389
GE.k | 5 | GE.k.5 | 1x1.25 | 2.1175
GE.k | 6 | GE.k.6 | 2x.63 | 1.2721
GE.l | 1 | GE.l.1 | 20 | 0.85607
GE.l | 2 | GE.l.2 | 15 | 0.93646
GE.l | 3 | GE.l.3 | 10 | 1
GE.l | 4 | GE.l.4 | 5 | 1.2582
GE.l | 5 | GE.l.5 | 1.25 | 1.3062
GE.m | 1 | GE.m.1 | 20 | 0.85707
GE.m | 2 | GE.m.2 | 15 | 0.95896
GE.m | 3 | GE.m.3 | 10 | 1
GE.m | 4 | GE.m.4 | 5 | 1.2006
GE.m | 5 | GE.m.5 | 1.25 | 1.1488
GE.n | 1 | GE.n.1 | 20 | 0.86018
GE.n | 2 | GE.n.2 | 15 | 0.93164
GE.n | 3 | GE.n.3 | 10 | 1
GE.n | 4 | GE.n.4 | 5 | 1.2352
GE.n | 5 | GE.n.5 | 2 x 0.63 | 1.277
GE.n | 6 | GE.n.6 | 1 x 1.25 | 2.4367
GE.o | 1 | GE.o.1 | 40 | 0.8582
GE.o | 2 | GE.o.2 | 20 | 0.92446
GE.o | 3 | GE.o.3 | 10 | 1.0989
GE.o | 4 | GE.o.4 | 1.25 | 1.4624
GE.p | 1 | GE.p.1 | 40 | 0.8582
GE.p | 2 | GE.p.2 | 20 | 0.92446
GE.p | 3 | GE.p.3 | 10 | 1.0989
GE.p | 4 | GE.p.4 | 1.25 | 1.4624
PH.e | 1 | PH.e.1 | 10 | 1
PH.e | 2 | PH.e.2 | 7 | 0.99
PH.e | 3 | PH.e.3 | 5 | 0.99942
PH.e | 4 | PH.e.4 | 3 | 0.98442
PH.e | 5 | PH.e.5 | 2 | 0.57
PH.e | 6 | PH.e.6 | 1.5 | 0.92035
PH.e | 7 | PH.e.7 | 1 | 3.0088
PH.e | 8 | PH.e.8 | 1 (AV) | 1.11
PH.f | 1 | PH.f.1 | 10 | 1
PH.f | 2 | PH.f.2 | 5 | 0.98319
PH.f | 3 | PH.f.3 | 2 | 1.1111
PH.i | 1 | PH.i.1 | 10 | 1
PH.i | 2 | PH.i.2 | 5 | 1
PH.i | 3 | PH.i.3 | 3 | 1
PH.i | 4 | PH.i.4 | 2 | 1.1
PH.k | 1 | PH.k.1 | 10 | 1
PH.k | 2 | PH.k.2 | 7 | 0.98479
PH.k | 3 | PH.k.3 | 5 | 0.9862
PH.k | 4 | PH.k.4 | 3 | 0.98269
PH.k | 5 | PH.k.5 | 2 | 0.64457
PH.k | 6 | PH.k.6 | 1 | 1.4491
PH.l | 1 | PH.l.1 | 20 | 0.89945
PH.l | 2 | PH.l.2 | 16 | 0.92929
PH.l | 3 | PH.l.3 | 10 | 1
PH.l | 4 | PH.l.4 | 4 | 1.2125
PH.l | 5 | PH.l.5 | 1 | 2.1891
PH.m | 1 | PH.m.1 | 8 | 1
PH.m | 2 | PH.m.2 | 5 | 0.95252
PH.m | 3 | PH.m.3 | 4 | 0.95324
PH.m | 4 | PH.m.4 | 3 | 0.9631
PH.m | 5 | PH.m.5 | 2 | 0.9813
PH.n | 1 | PH.n.1 | 24 (8 x 3) | 0.87861
PH.n | 2 | PH.n.2 | 24 (16 x 1.5) | 0.89731
PH.n | 3 | PH.n.3 | 12 (4 x 3) | 0.92814
PH.n | 4 | PH.n.4 | 12 (16 x 0.75) | 1
PH.n | 5 | PH.n.5 | 6 | 1.0986
PH.n | 6 | PH.n.6 | 3 | 1.5107
PH.o | 1 | PH.o.1 | 10 | 1
PH.o | 2 | PH.o.2 | 7 | 1.011
PH.o | 3 | PH.o.3 | 5 | 1.009
PH.o | 4 | PH.o.4 | 3 | 1.0421
PH.o | 5 | PH.o.5 | 2 | 1.0423
PH.o | 6 | PH.o.6 | 1 | 2.0847
PH.p | 1 | PH.p.1 | 24 (8 x 3) | 0.86387
PH.p | 2 | PH.p.2 | 24 (16 x 1.5) | 0.88445
PH.p | 3 | PH.p.3 | 18 | 0.8997
PH.p | 4 | PH.p.4 | 12 (4 x 3) | 0.93009
PH.p | 5 | PH.p.5 | 12 (16 x 0.75) | 1
PH.p | 6 | PH.p.6 | 6 | 1.1465
PH.p | 7 | PH.p.7 | 3 | 1.6398
PH.p | 8 | PH.p.8 | 1.2 | 1.3887
PH.q | 1 | PH.q.1 | 40 | 0.75
PH.q | 2 | PH.q.2 | 25 | 0.88
PH.q | 3 | PH.q.3 | 15 | 0.87
PH.q | 4 | PH.q.4 | 10 | 1
PH.q | 5 | PH.q.5 | 7.5 | 1.02
PH.q | 6 | PH.q.6 | 1.25 | 1.89
PH.q | 7 | PH.q.7 | 1 | 1.9
PI.b | 1 | PI.b.1 | 10 | 1
PI.b | 2 | PI.b.2 | 5 | 0.95
PI.b | 3 | PI.b.3 | 3 | 0.95
PI.b | 4 | PI.b.4 | 1.5 | 0.92
PI.c | 1 | PI.c.1 | 10 | 1
PI.c | 2 | PI.c.2 | 8 | 0.94712
PI.c | 3 | PI.c.3 | 5 | 0.94324
PI.c | 4 | PI.c.4 | 4 | 0.93088
PI.c | 5 | PI.c.5 | 3 | 0.92457
PI.c | 6 | PI.c.6 | 2 | 0.92394
PI.c | 7 | PI.c.7 | 1.5 | 0.91322
PI.c | 8 | PI.c.8 | 1 | 0.8663
PI.d | 1 | PI.d.1 | 20 | 0.89945
PI.d | 2 | PI.d.2 | 16 | 0.92929
PI.d | 3 | PI.d.3 | 10 | 1
PI.d | 4 | PI.d.4 | 4 | 1.2125
PI.d | 5 | PI.d.5 | 1 | 2.1891
PI.e | 1 | PI.e.1 | 8 | 1
PI.e | 2 | PI.e.2 | 5 | 0.95252
PI.e | 3 | PI.e.3 | 4 | 0.95324
PI.e | 4 | PI.e.4 | 3 | 0.9631
PI.e | 5 | PI.e.5 | 2 | 0.9813
SI.d | 1 | SI.d.1 | 10 | 1
SI.d | 2 | SI.d.2 | 8 | 1
SI.d | 3 | SI.d.3 | 5 | 1
SI.d | 4 | SI.d.4 | 3 | 0.94
SI.d | 5 | SI.d.5 | 2 | 0.95
SI.d | 6 | SI.d.6 | 1 | 1.9
SI.e | 1 | SI.e.1 | 10 | 1
SI.e | 2 | SI.e.2 | 5 | 0.96716
SI.e | 3 | SI.e.3 | 3 | 0.97612
SI.e | 4 | SI.e.4 | 2 | 0.92537
SI.e | 5 | SI.e.5 | 1 | 1.8448
SI.g | 1 | SI.g.1 | 10 | 1
SI.g | 2 | SI.g.2 | 5 | 0.90541
SI.g | 3 | SI.g.3 | 3 | 1
SI.g | 4 | SI.g.4 | 2 | 0.75676
SI.g | 5 | SI.g.5 | 1 | 1.3378
SI.h | 1 | SI.h.1 | 10 | 1
SI.h | 2 | SI.h.2 | 5 | 0.88288
SI.h | 3 | SI.h.3 | 2 | 0.98198
SI.h | 4 | SI.h.4 | 1 | 0.94595
SI.i | 1 | SI.i.1 | 10 | 1
SI.i | 2 | SI.i.2 | 8 | 1.0136
SI.i | 3 | SI.i.3 | 5 | 1.0243
SI.i | 4 | SI.i.4 | 3 | 0.94929
SI.i | 5 | SI.i.5 | 2 | 0.98426
SI.i | 6 | SI.i.6 | 1 | 1.2575
SI.j | 1 | SI.j.1 | 20 (Zoom) | 0.91726
SI.j | 2 | SI.j.2 | 20 (Access) | 0.97442
SI.j | 3 | SI.j.3 | 16 (Zoom) | 0.88228
SI.j | 4 | SI.j.4 | 16 (Access) | 0.97629
SI.j | 5 | SI.j.5 | 10 | 1
SI.j | 6 | SI.j.6 | 5 | 1.0413
SI.j | 7 | SI.j.7 | 4 | 1.2293
SI.j | 8 | SI.j.8 | 2 | 1.2973
SI.j | 9 | SI.j.9 | 1 (Zoom) | 2.0837
SI.j | A | SI.j.A | 1 (Access) | 2.1245
SI.k | 1 | SI.k.1 | 10 | 1
SI.k | 2 | SI.k.2 | 8 | 1.012
SI.k | 3 | SI.k.3 | 5 | 1.0014
SI.k | 4 | SI.k.4 | 3 | 1.0149
SI.k | 5 | SI.k.5 | 2 | 0.97448
SI.l | 1 | SI.l.1 | 20 | 0.91601
SI.l | 2 | SI.l.2 | 16 | 0.80461
SI.l | 3 | SI.l.3 | 10 | 1.0004
SI.l | 4 | SI.l.4 | 4, lrg fs | 1.1993
SI.l | 5 | SI.l.5 | 4, sm fs | 1.1665
SI.l | 6 | SI.l.6 | 1 | 2.0303
SI.m | 1 | SI.m.1 | 24 | 1.1146
SI.m | 2 | SI.m.2 | 18 | 1.1342
SI.m | 3 | SI.m.3 | 12 | 1.2114
SI.m | 4 | SI.m.4 | 10 | 1
SI.m | 5 | SI.m.5 | 9 | 1.2877
SI.m | 6 | SI.m.6 | 2 | 1.1026
SI.m | 7 | SI.m.7 | 1.2 | 1.554
SI.n | 1 | SI.n.1 | 24 | 1.1146
SI.n | 2 | SI.n.2 | 18 | 1.1714
SI.n | 3 | SI.n.3 | 12 | 1.2855
SI.n | 4 | SI.n.4 | 10 | 1.01
SI.n | 5 | SI.n.5 | 9 | 1.3784
SI.n | 6 | SI.n.6 | 2 | 1.1154
SI.n | 7 | SI.n.7 | 1.2 | 1.5728
SI.o | 1 | SI.o.1 | 1 | 0.99502
SI.o | 2 | SI.o.2 | 3 | 1.3304
SI.o | 3 | SI.o.3 | 6 | 1.3479
SI.o | 4 | SI.o.4 | 10 | 1
SI.o | 5 | SI.o.5 | 12 | 1.1501
SI.o | 6 | SI.o.6 | 18 | 1.0811
SI.p | 1 | SI.p.1 | 1.2 | 1.5933
SI.p | 2 | SI.p.2 | 2 | 1.4054
SI.p | 3 | SI.p.3 | 6 | 1.4851
SI.p | 4 | SI.p.4 | 7.5 | 1.4275
SI.p | 5 | SI.p.5 | 10 | 1
SI.p | 6 | SI.p.6 | 12 | 1.2431
SI.p | 7 | SI.p.7 | 15 | 1.2059
SI.p | 8 | SI.p.8 | 18 | 1.1565
SI.q | 1 | SI.q.1 | 5 | 1.0027
SI.q | 2 | SI.q.2 | 10 | 1
SI.q | 3 | SI.q.3 | 6 (20* x 0.6) | 1.121
SI.q | 4 | SI.q.4 | 12 (20 x 0.6) | 1.3647
SI.q | 5 | SI.q.5 | 12 (40* x 0.6) | 1.2582
SI.q | 6 | SI.q.6 | 28.8 | 1.1195
SI.r | 1 | SI.r.1 | 2 | 1.4531
SI.r | 2 | SI.r.2 | 5 | 0.99198
SI.r | 3 | SI.r.3 | 10 | 1
SI.r | 4 | SI.r.4 | 3.6 (12* x 0.6) | 1.9391
SI.r | 5 | SI.r.5 | 6 (20* x 0.6) | 1.0522
SI.r | 6 | SI.r.6 | 18 | 1.1868
SI.r | 7 | SI.r.7 | 19.2 (64* x 0.6) | 1.2105
SI.r | 8 | SI.r.8 | 28.8 | 1.1214
SI.s | 1 | SI.s.1 | 2 | 1.6653
SI.s | 2 | SI.s.2 | 5 | 1.0261
SI.s | 3 | SI.s.3 | 6 (20* x 0.6) | 1.4476
SI.s | 4 | SI.s.4 | 10 | 1
SI.s | 5 | SI.s.5 | 12 (40* x 0.6) | 1.4768
SI.s | 6 | SI.s.6 | 14.4 | 1.3894
SI.s | 7 | SI.s.7 | 18 | 1.3324
SI.s | 8 | SI.s.8 | 19.2 (16 x 1.2) | 1.3226
SI.s | 9 | SI.s.9 | 19.2 (64* x 0.6) | 1.3627
SI.s | A | SI.s.A | 19.2 (32 x 0.6) | 1.3226
TO.b | 1 | TO.b.1 | 10 | 1
TO.b | 2 | TO.b.2 | 5 | 0.8866
TO.b | 3 | TO.b.3 | 2 | 0.84536
TO.b | 4 | TO.b.4 | 1 | 1.5258
TO.e | 1 | TO.e.1 | 10 | 1
TO.e | 2 | TO.e.2 | 5 | 1.02
TO.e | 3 | TO.e.3 | 3 | 1.15
TO.e | 4 | TO.e.4 | 2 | 1.37
TO.e | 5 | TO.e.5 | 1 | 2.75
TO.g | 1 | TO.g.1 | 10 | 1
TO.g | 2 | TO.g.2 | 7 | 1
TO.g | 3 | TO.g.3 | 5 | 0.99
TO.g | 4 | TO.g.4 | 3 | 1.1
TO.g | 5 | TO.g.5 | 2 | 1.22
TO.g | 6 | TO.g.6 | 1 | 1.84
TO.h | 1 | TO.h.1 | 32 | 0.87346
TO.h | 2 | TO.h.2 | 20 | 0.91347
TO.h | 3 | TO.h.3 | 16 | 0.93086
TO.h | 4 | TO.h.4 | 12 | 0.98158
TO.h | 5 | TO.h.5 | 8 | 1
TO.h | 6 | TO.h.6 | 4 | 1.3709
TO.h | 7 | TO.h.7 | 2 | 2.0736
TO.i | 1 | TO.i.1 | 32 | 0.77453
TO.i | 2 | TO.i.2 | 20 | 0.83695
TO.i | 3 | TO.i.3 | 16 | 0.8674
TO.i | 4 | TO.i.4 | 12 | 0.91864
TO.i | 5 | TO.i.5 | 8 | 1
TO.i | 6 | TO.i.6 | 4 | 1.3544
TO.i | 7 | TO.i.7 | 2 | 2.0747
TO.k | 1 | TO.k.1 | 32 | 0.86124
TO.k | 2 | TO.k.2 | 20 | 0.88527
TO.k | 3 | TO.k.3 | 16 | 0.90476
TO.k | 4 | TO.k.4 | 12 | 0.97966
TO.k | 5 | TO.k.5 | 8 | 1
TO.k | 6 | TO.k.6 | 4 | 1.3216
TO.k | 7 | TO.k.7 | 2 | 1.9689
TO.l | 1 | TO.l.1 | 20 | 0.89555
TO.l | 2 | TO.l.2 | 16 | 0.92184
TO.l | 3 | TO.l.3 | 10 | 1
TO.l | 4 | TO.l.4 | 8 | 1.0725
TO.l | 5 | TO.l.5 | 6 | 1.2168
TO.l | 6 | TO.l.6 | 4 | 1.2859
TO.l | 7 | TO.l.7 | 2 | 2.009
TO.l | 8 | TO.l.8 | 1 | 3.3547
TO.m | 1 | TO.m.1 | 32 (4 x 8) | 0.83766
TO.m | 2 | TO.m.2 | 32 (16 x 2) | 0.84596
TO.m | 3 | TO.m.3 | 24 | 0.83407
TO.m | 4 | TO.m.4 | 16 (4 x 4) | 0.86276
TO.m | 5 | TO.m.5 | 16 (16 x 1) | 0.93079
TO.m | 6 | TO.m.6 | 12 | 0.8481
TO.m | 7 | TO.m.7 | 8 (1 x 8) | 0.92888
TO.m | 8 | TO.m.8 | 8 (4 x 2) | 1
TO.m | 9 | TO.m.9 | 8 (16 x 0.5) | 1.1007
TO.m | A | TO.m.A | 6 | 0.94051
TO.m | B | TO.m.B | 4 (1 x 4) | 1.2331
TO.m | C | TO.m.C | 4 (4 x 1) | 1.3849
TO.m | D | TO.m.D | 2 (1 x 2) | 1.9316
TO.m | E | TO.m.E | 2 (4 x 0.5) | 2.0734
TO.m | F | TO.m.F | 1 | 3.2783
Note on D236: Uses z-axis flying focal spot to double number of data sets per revolution
Note on D238: Uses z-axis flying focal spot to double number of data sets per revolution
Note on E35: Value taken from LightSpeed Ultra data

## Doses
Slab | No. of Organs | Adrenals | Brain | Breasts | Eye Lenses | Gall Bladder | Stomach | Small Intestine | Upper Large Intestine | Lower Large Intestine | Heart | Kidneys | Liver | Lungs | Ovaries | Pancreas | Skin | Spleen | Testicles | Thymus | Thyroid | Urinary Bladder | Uterus | Head Region | Trunk Region | Leg Region | Skeleton | Red Marrow
1
2
3
4
5
6
7
8
9
10
11
12
13
14
15
16
17
18
19
20
21
22
23
24
25
26
27
28
29
30
31
32
33
34
35
36
37
38
39
40
41
42
43
44
45
46
47
48
49
50
51
52
53
54
55
56
57
58
59
60
61
62
63
64
65
66
67
68
69
70
71
72
73
74
75
76
77
78
79
80
81
82
83
84
85
86
87
88
89
90
91
92
93
94
95
96
97
98
99
100
101
102
103
104
105
106
107
108
109
110
111
112
113
114
115
116
117
118
119
120
121
122
123
124
125
126
127
128
129
130
131
132
133
134
135
136
137
138
139
140
141
142
143
144
145
146
147
148
149
150
151
152
153
154
155
156
157
158
159
160
161
162
163
164
165
166
167
168
169
170
171
172
173
174
175
176
177
178
179
180
181
182
183
184
185
186
187
188
189
190
191
192
193
194
195
196
197
198
199
200
201
202
203
204
205
206
207
208

## DoseCalculations
Doses to organs in the phantom
| Adrenals | Brain | Breasts | Eye Lenses | Gall Bladder | Stomach | Small Intestine | Upper Large Intestine | Lower Large Intestine
SvDf | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
HT | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Heart | Kidneys | Liver | Lungs | Ovaries | Pancreas | Skin | Spleen | Testicles
SvDf | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
HT | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Thymus | Thyroid | Urinary Bladder | Uterus | Head Region | Trunk Region | Leg Region | Skeleton | Red Marrow
SvDf | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
HT | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Mass weighted doses to organs (used for calculation of remainder doses)
| Adrenals | Brain | Breasts | Eye Lenses | Gall Bladder | Stomach | Small Intestine | Upper Large Intestine | Lower Large Intestine
HT | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Organ mass | 14 | 1400 | 301.5 | 0 | 10 | 150 | 640 | 209 | 158
Mass wt Dose | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Heart | Kidneys | Liver | Lungs | Ovaries | Pancreas | Skin | Spleen | Testicles
HT | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Organ mass | 570 | 310 | 1810 | 999 | 8.27 | 100 | 2862 | 180 | 37.1
Mass wt Dose | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Thymus | Thyroid | Urinary Bladder | Uterus | Head Region | Trunk Region | Leg Region | Skeleton | Red Marrow
HT | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Organ mass | 20 | 19.6 | 45.1 | 80 | 5769 | 43048 | 22093 | 9516 | 1120
Mass wt Dose | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Oesophagus | Stomach Contents | SI Contents | ULI Contents | LLI Contents | Heart Contents | Bladder Contents | Oesoph. Contents | Whole Body | Muscle
HT | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Organ mass | 90 | 200 | 400 | 220 | 135 | 260 | 102 | 50 | 70910 | 48893.43
Mass wt Dose | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
Remainder organs
Max? | Organ | HT | Organ mass | Mass wt Dose | HT (2 significant figures) | ICRP 60 organ | ICRP 103 organ
#N/A | Thymus | #N/A | 20 | #N/A | #N/A | Yes | Yes
#N/A | Adrenals | #N/A | 14 | #N/A | #N/A | Yes | Yes
#N/A | Pancreas | #N/A | 100 | #N/A | #N/A | Yes | Yes
#N/A | Spleen | #N/A | 180 | #N/A | #N/A | Yes | Yes
#N/A | Kidneys | #N/A | 310 | #N/A | #N/A | Yes | Yes
#N/A | Small Int | #N/A | 640 | #N/A | #N/A | Yes | Yes
#N/A | Uterus | #N/A | 80 | #N/A | #N/A | Yes | Yes
#N/A | Muscle | #N/A | 28000 | #N/A | #N/A | Yes | Yes
#N/A | Brain | #N/A | 1400 | #N/A | #N/A | Yes | No
- | Gall Bladder | #N/A | N/A | N/A | #N/A | No | Yes
- | Heart | #N/A | N/A | N/A | #N/A | No | Yes
- | ET region (Thyroid) | #N/A | N/A | N/A | #N/A | No | Yes
- | Lymph nodes (Muscle) | #N/A | N/A | N/A | #N/A | No | Yes
- | Oral mucosa (Brain) | #N/A | N/A | N/A | #N/A | No | Tes
ICRP 60 Remainder calculations: Remainder dose is mass weighted mean of remianders, unless Max(remainder) > Max(organs with weighting factors)
where remainder is split in two, first is Max(remainder), second is mean of remainders w/o Max(remainder)
Remainder weighted mean dose (ICRP60) |  |  |  | #N/A
Maximum Remainder dose |  |  | #N/A | #N/A
Use Max remainder organ? |  |  |  | #N/A
Remainder mean w/o max dose |  |  |  | #N/A
ICRP 103 Remainder calculations: Remainder dose is mean of remainder organs
Remainder mean dose (ICRP103) |  |  |  | #N/A
Organs with ICRP 60 / 103 Weighting factors
Organ |  | ICRP60 wT | ICRP103 wT | Selected wT | HT | HT (2 significant figures) | wT.HT | wT.HT (2 significant figures)
Gonads |  | 0.2 | 0.08 | 0.2 | #N/A | #N/A | #N/A | #N/A
Bone Marrow (red) |  | 0.12 | 0.12 | 0.12 | #N/A | #N/A | #N/A | #N/A
Colon |  | 0.12 | 0.12 | 0.12 | #N/A | #N/A | #N/A | #N/A
Lung |  | 0.12 | 0.12 | 0.12 | #N/A | #N/A | #N/A | #N/A
Stomach |  | 0.12 | 0.12 | 0.12 | #N/A | #N/A | #N/A | #N/A
Bladder |  | 0.05 | 0.04 | 0.05 | #N/A | #N/A | #N/A | #N/A
Breast |  | 0.05 | 0.12 | 0.05 | #N/A | #N/A | #N/A | #N/A
Liver |  | 0.05 | 0.04 | 0.05 | #N/A | #N/A | #N/A | #N/A
Oesophagus (Thymus) |  | 0.05 | 0.04 | 0.05 | #N/A | #N/A | #N/A | #N/A
Thyroid |  | 0.05 | 0.04 | 0.05 | #N/A | #N/A | #N/A | #N/A
Skin |  | 0.01 | 0.01 | 0.01 | #N/A | #N/A | #N/A | #N/A
Bone Surface |  | 0.01 | 0.01 | 0.01 | #N/A | #N/A | #N/A | #N/A
Brain |  | 0 | 0.01 | 0 | #N/A | #N/A | #N/A | #N/A
Salivary Glands (Brain) |  | 0 | 0.01 | 0 | #N/A | #N/A | #N/A | #N/A
Remainder |  | #N/A | 0.12 | #N/A | #N/A | #N/A | #N/A | #N/A
#N/A |  | #N/A | 0 | #N/A | #N/A | #N/A | #N/A | #N/A
[6 empty] | Total Effective Dose (mSv) | #N/A | #N/A
Other organs
|  | HT | HT (2 significant figures)
Eye lenses |  | #N/A | #N/A
Testes |  | #N/A | #N/A
Ovaries |  | #N/A | #N/A
Uterus |  | #N/A | #N/A
Prostate (bladder) |  | #N/A | #N/A
Note on A3: Sum of normalised organ doses from all slices along scan length
Note on A4: Organ dose (mGy)
Note on A13: Organ dose (mGy)
Note on A14: SR250 organ masses. In cells with a sum of two values, first is mass of organ, second is mass of contents
Note on A15: Organ dose weighted according to organ mass (used in muscle dose calculation)
Note on B42: SR250 does not include the extrathoracic region listed in ICRP103, closest approximation is thyroid
Note on B43: SR250 does not include the lymph nodes as listed in ICRP103, closest approximation is muscle
Note on B44: SR250 does not include the oral mucosa listed in ICRP103, closest approximation is brain
Note on B69: SR250 does not include salivary gland, the closest approximation is brain
Note on B71: This row only contains an organ if the maximum ICRP60 remainder organ dose is greater than the maximum dose to organs with specified weighting factors
Note on C76: Organ Dose (mGy)
Note on E55: Organ Tissue weighting factor
Note on F55: Organ Dose (mGy)
Note on H55: Effective dose (Organ dose x Tissue weighting factor) rounded to two significant digits
Note on J26: Sum of Head, Trunk and Leg Region masses
Note on K26: Whole Body mass - (organ mass + organ contents mass)

## Selections
Manufacturer: | 7
1 | CGR | CG
2 | Elscint | EL
3 | GE | GE
4 | Philips | PH
5 | Picker / Marconi | PI
6 | Shimadzu | SH
7 | Siemens | SI
8 | Toshiba | TO
| Current Selection: | SI
Scanner: | 19
1 | Siemens CR, CR512, DRH | SI.a
2 | Siemens Somatom 2, DR1/2/3 | SI.b
3 | Siemens DRG, DRG1 | SI.c
4 | Siemens Somatom Plus 4 Series | SI.d
5 | Siemens Somatom AR-C, AR.SP, AR-T | SI.e
6 | Siemens AR.HP | SI.f
7 | Siemens Plus, DXP, Plus-S | SI.g
8 | Siemens Hi Q | SI.h
9 | Siemens Balance, Emotion | SI.i
10 | Siemens Volume Zoom, Access | SI.j
11 | Siemens Emotion Duo | SI.k
12 | Siemens Sensation 4 | SI.l
13 | Siemens Sensation 16 | SI.m
14 | Siemens Sensation 16 Straton | SI.n
15 | Siemens Emotion 6 | SI.o
16 | Siemens Sensation 10 | SI.p
17 | Siemens Sensation Open | SI.q
18 | Siemens Sensation 64 | SI.r
19 | Siemens Definition AS | SI.s
| #N/A | SI.s
kV | 0
1 | 080 | SI.s.080
2 | 100 | SI.s.100
3 | 120 | SI.s.120
4 | 140 | SI.s.140
5 | #N/A | TO.a.120
| #N/A | 0
Body Region | 1
1 | Head
2 | Body
| Current Selection: | Head
Match Set: | #N/A
Match: | #N/A
Current Data: | 0
Data set is: | #N/A
Phantom Slabs: | 106 | to | 149
Whole Slabs: | 106 | to | 148
Start Slab: | 0 | of slab | 106
End Slab: | 0 | of slab | 149
Collimation | 3
1 | 2 | SI.s.1
2 | 5 | SI.s.2
3 | 6 (20* x 0.6) | SI.s.3
4 | 10 | SI.s.4
5 | 12 (40* x 0.6) | SI.s.5
6 | 14.4 | SI.s.6
7 | 18 | SI.s.7
8 | 19.2 (16 x 1.2) | SI.s.8
9 | 19.2 (64* x 0.6) | SI.s.9
A | 19.2 (32 x 0.6) | SI.s.A
B | #N/A | SI.s.B
C | #N/A | SI.s.C
D | #N/A | SI.s.D
E | #N/A | SI.s.E
F | #N/A | SI.s.F
| #N/A | SI.s.3
1.4476
6 (20* x 0.6)mm
ICRP OWS | 1
1 | ICRP 60
2 | ICRP 103
Current value | ICRP 60
Note on B54: Test to see if start slab is the same as the end slab, if so, subtract fraction of slab up to the start point from the end slab fraction